Biocompatibility

- The following describes the human body contacting materials of {填写申报产品名称 } (model: {填写申报产品的具体型号}).

{填写描述与人体直接或间接接触的方式，以及对应的材料信息}

（参考示例：components of the device contacting with human is PE film. The material used for this components is Polyethylene PE (100%), which contact with human body only in short-term period. The different models of the device is made with the same material. They have the same intended use. The only different between them are size. So we consider the test result of Model 1 can cover other models.）

Therefore, pursuant to the standard ISO10993-1:2009 biological evaluation of medical devices - part 1: evaluation and testing, the human body contacting material require the following biocompatibility tests:

- ISO10993-5: 2009 biological evaluation of medical devices-part 5: Tests for in vitro Cytotoxicity.

- ISO10993-10:2010 biological evaluation of medical devices-part 10: Tests for Irritation and skin sensitization.

【生物相容性涉及的标准，需要结合产品实际情况，根据ISO 10993-1进行判断，并得出具体需要测试的项目，并按照相应的产品进行测试，同时需要考虑指南

“Use of International Standard ISO 10993-1, "Biological evaluation of medical devices - Part 1: Evaluation and testing within a risk management process.”的要求。】

The human contacting material has passed all of the above testings.

The list of all patient contacting components as follows.

{填写与人体直接或间接接触的材料的信息}

（参考示例：

| Device Components | Contact with human body | Contact Method | Contact classification |  | Tissue-contacting device component |  |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Category | Contact Duration | Materials | Additive |
| Film | Oral cavity and axillary cavity | Direct contact | Surface contacting | Less than 24 hour duration | PE(100%) | Do not contains additive |

- ）

【如果产品还有颜色添加剂，则需要明确颜色添加剂的成分，颜色添加剂需要符合CFR中颜色添加剂的相关要求】

Biocompatibility

{概述生物相容性相关信息}

（参考示例：

The device contains component that come into direct contact with patients, and the patient-contacting component is PE film. The contacting materials is Polyethylene PE (100%), no color additives contained.）

Based on the standard ISO 10993-1:2009 biological evaluation of medical devices -Part 1: evaluation and testing, the {填写申报产品名称 } is tested by the third institution which is {填写生物相容性测试机构 } and Summary of test as follow:

| Device Name | {填写测试产品名称 } |
| --- | --- |
| Manufacturer | {填写测试产生产企业 } |
| Contacting Components | Component name: {填写与人体接触部分的描述 } Component material: {填写与人体接触部分的材质} |
| Contact classification per ISO 10993-1 | Nature of body contact: Category: Surface device Contact: Oral cavity Contact duration: limited(≤24) |
| Test items | - Test for in vitro cytotoxicity - Test for skin sensitization - Test for oral mucosa Irritation Test |
| Test standards | - ISO 10993-5:2009 Biological evaluation of medical devices- Part 5: Tests for in vitro cytotoxicity 2) ISO 10993-10:2010 Biological evaluation of medical devices- Part 10: Tests for Irritation and skin sensitization |
| Test result | {填写5份测试报告中的结论 } 【在FDA做生物相容性测试，要求按照GLP流程进行测试，同时按照双极性进行试验】 |
| - Test report No. | {填写5份测试报告中报告编号 } |

【以上以最常规的与皮肤接触为例，罗列相应的涉及到的测试信息】

The results of test items above passed, proving that using {填写测试产品名称 } which manufactured by {填写测试产生产企业 } is safety for biocompatibility.

Detailed Test Report as below:

Appendix 1 In Vitro Cytotoxicity Test

Appendix 2 Skin Sensitization Test(Sodium Chloride Injection Extract)

Appendix 3 Skin Sensitization Test(Sesame oil Extract)

Appendix 4 Oral mucosa Irritation Test(Sodium Chloride Injection Extract)

Appendix 5 Oral mucosa Irritation Test(Sesame oil Extract)

.........

【将涉及到的生物相容性报告分别附在后面】