# Guidance for the Content of Premarket Notifications for Biopsy Devices Used in Gastroenterology and Urology (Text Only)

- Share

- Tweet

- Linkedin

- Pin it

- More sharing options

  - Linkedin

  - Pin it

- Email

- Print

This guidance was written prior to the February 27, 1997 implementation of FDA’s Good Guidance Practices, GGP’s. It does not create or confer rights for or on any person and does not operate to bind FDA or the public. An alternative approach may be used if such approach satisfies the requirements of the applicable statute, regulations, or both. This guidance will be updated in the next revision to include the standard elements of GGP’s.

Biopsy devices for gastroenterology and urology are described in the FDA

regulations under three classifications.  These classifications are 21 CFR

876.1075 (a) Gastroenterology-Urology Biopsy Instruments, 21 CFR 876.1500

(a) Endoscopes and Accessories, and 21 CFR 876.4300 (a) Endoscopic

Electrosurgical Unit and Accessories.

The classification 876.1075 (a), Gastroenterology-Urology Biopsy

Instruments, includes devices that are "used to remove, by cutting or

aspiration, a specimen of tissue for microscopic examination.  This generic

type of device includes the biopsy punch, gastrointestinal mechanical

biopsy instrument, suction biopsy instrument, gastro-urology biopsy needle

and needle set, and nonelectric biopsy forceps.  This section does not

apply to biopsy instruments covered by classification regulations in other

parts of the device classification regulations."  These are Class II

devices as stated in 21 CFR 876.1075 (b) and their procodes are:  78 FFF -

cover, biopsy forceps; 78 FCL - forceps, biopsy, non-electric; 78 KNW -

instrument, biopsy; 78 FCF - instrument, biopsy, mechanical,

gastrointestinal; 78 FCK - instrument, biopsy, suction; 78 FCI - punch,

biopsy; and 78 FCG - set, biopsy needle and needle, gastro-urology.

The classification 876.1500 (a), Endoscopes and Accessories, includes a

"device ... to allow observation and manipulation of body cavities, hollow

organs, and canals."  Cytology brushes, as endoscopic accessories, are

covered under this classification, are Class II devices as stated in 21 CFR

876.1500 (b), and have the procode 78 FDX - brush, cytology, for

endoscopes.

The classification 876.4300, Endoscopic Electrosurgical Unit and

Accessories, includes a "device used to perform electrosurgical procedures

through an endoscope.  This generic type of device includes ... electric

biopsy forceps, ... flexible snare, ... self-opening rigid snare, ... "

These are Class II devices as stated in 21 CFR 876.4300 (b) and their

procodes are 78 KGE - forceps, biopsy, electric; 78 FDI - snare, flexible;

and 78 FDJ - snare, rigid self-opening.

The primary reference for the information required to be in a premarket

notification (510(k)) for a medical device is set forth in 21 CFR 807.87.

The purpose of this regulation is to provide adequate documented

information to determine substantial equivalence to a device in commercial

distribution.  Substantial equivalence is to be established with respect

to, but not limited to, intended use, design, energy used/delivered,

materials, performance, safety, effectiveness, labeling, and other

applicable characteristics.

FDA recommends that each premarket notification for a biopsy device include

the following information in order to ensure that the submission is

complete and will permit a determination of substantial equivalence:

I.  The device name, including both the trade or proprietary name and the

classification name (Gastroenterology-Urology Biopsy Instrument,

Endoscopes and Accessories for cytology brushes, or Endoscopic

Electrosurgical Unit and Accessories for electric biopsy forceps) of

the device as described in 21 CFR 807.87 (a).

II. The establishment registration number, if applicable, of the owner or

operator submitting the premarket notification as described in 21 CFR

807.87 (b).

III.The class (Class II) in which the device has been placed under section

513 of the act and the appropriate panel (78 Gastroenterology/Urology)

as described in 21 CFR 807.87 (c).

IV. Action taken by the person required to register to comply with the

requirements of the act under section 513 for Special Controls.  Note

that Special Controls are not currently required for biopsy devices

under section 513 of the act.

V.  The Safe Medical Devices Act of 1990 (SMDA) requires all persons

submitting a premarket notification submission to include either (1) a

summary of the safety and effectiveness information in the premarket

notification submission upon which an equivalence determination could

be based (510(k) summary), OR (2) a statement that safety and

effectiveness information will be made available to interested persons

upon request (510(k) statement).  Safety and effectiveness information

refers to information in the premarket notification submission,

including adverse safety and effectiveness information, that is

relevant to an assessment of substantial equivalence.  The information

could be descriptive information about the new and predicate device(s),

or performance or clinical testing information.

VI. Proposed labels, labeling, and advertisements sufficient to describe

the biopsy device, its intended use, and the directions for use should

be provided with a specific intended use statement and any warnings,

contraindications, or limitations clearly displayed as described in 21

CFR 807.87 (e).  The label of the device must bear the caution

statement as outlined in 21 CFR 801.109 (b) (1):  "CAUTION:  Federal

law restricts this device to sale by or on the order of a physician."

A.  A label includes any identification on the biopsy device and on

the package in which it is stored and shipped.  The  package

device label should include the device name, U.S. point of

contact, corporation name, address, and phone number.  The package

label should also include all of the above, as well as sterility

status, expiration date, disposable/single use, quantity enclosed,

size (diameter and length), energy used/delivered, etc.

B.  Device labeling for the biopsy device includes the intended use, a

description of the device, and directions for use.

1.  The intended use statement should include specific

indications.  The anatomical sites, i.e. target organs, and the

target population should be defined.

2.  The directions for use should contain comprehensive

instructions to include, but not necessarily be limited to, how to

prepare the biopsy device for use, how to operate the biopsy

device, how to stop operation and retrieve the sample, any special

instructions for specimen processing, a description of which parts

are single use/disposable or reusable, and a description of the

functional test procedures which should be performed for the

biopsy device prior to use.  Maintenance and troubleshooting

procedures should be outlined with instructions on how to perform

the maintenance, the frequency of the maintenance, and a

corporation contact point if troubleshooting procedures fail.

3.  Contraindications (e.g. excluded population), precautions

(e.g. needle advancement/penetration depth), and warnings should

be included in the labeling of the device.

C.  Advertisements or promotional literature for the biopsy device

that will accompany the device should be provided.  Literature or

labeling may not imply approval by FDA in any manner.  Guidance on

labeling issues is described in Bluebook Memo G91-1 "Device

Labeling Guidance (3/8/91)" and a copy may be obtained from the

Center for Devices and Radiological Health's Division of Small

Manufacturers Assistance at (800) 638-3041 or (301) 443-6597.

VII. A Summary of Equivalence comparing similar devices legally in

commercial distribution in the United States must be provided.  This

includes devices in commercial distribution prior to May 28, 1976, the

enactment date of the Medical Devices Amendments, and any new devices

introduced subsequently.  A Summary of Equivalence includes similarities

and differences between the device and the device to which it is compared.

The biopsy device should be compared with a legally marketed biopsy device,

including, but not limited to, the following:

intended use including specific target organs;

design

needle, cannula(s), and stylet size:  diameter and

gauge;

needle advancement/penetration depth;

sample notch size;

number of samples;

jaw size;

brush size: diameter and length;

mechanics of action (automatic, manual, spring

operated, electrically operated, etc.);

mode of action (single puncture and sample, single

puncture and multiple samples, brush rotation,

suction, scraping: side-to-side or up-and-down,

etc.);

electrical compatibility;

safety features; and

other applicable specifications

energy used/delivered for electric biopsy forceps and

biopsy guns;

materials of all components identifying those that come

into patient contact;

performance;

anatomical sites (i.e. target organs);

target population justifying any new population cited;

visualization techniques;

method of placement; and

other related information.

State whether the substantially equivalent device is a pre-amendment device

or a device which has been through the 510(k) process, providing the 510(k)

document control number if known.  The summary of equivalence information

should be provided in a manner that is clear and comprehensible, e.g.

tabular form.

VIII. For a device that has undergone a change or modification that could

significantly affect the safety or effectiveness of the device, or the

device is to be marketed for a new or different indication for use, the

510(k) must include appropriate supporting data to show that the

manufacturer has considered what consequences and effects the change or

modification or new use might have on the safety and effectiveness of the

device, as described in 21 CFR 807.87 (g).

Significant modifications should be supported by a rationale for the

modification with supporting documentation, including clinical or other

valid scientific studies which demonstrate that these differences do not

affect safety and effectiveness, as described in 21 CFR 807.87 (f).

The description of all biopsy devices should include any significant

changes or modifications from the predicate device that could affect

safety, effectiveness, or intended use.  Provide any bench, animal,

clinical, functional, in vitro, and/or any other testing data to support

your claims.  Provide certification regarding any compliance with voluntary

standards, if applicable.

IX.  The physical description of each biopsy device to be marketed should

be provided in the form of a labeled diagram, photograph/picture,

schematic, etc., which includes all internal/external,

assembled/unassembled, etc. parts of the biopsy device.  The physical

description should include the specifications (length, width, height,

diameter, power requirements, other applicable information) of the biopsy

device.  The description should also identify any parts which are

disposable (i.e., needles, sheath, cannula, etc.).  The labeled diagram,

photograph/picture, schematic, etc., should address the name and function

of all parts of the biopsy device.  This is especially important for

automated biopsy guns and electric biopsy forceps where part of the

mechanics of action is performed without operator intervention.

If the biopsy device is sold in a set that includes accessories, these

accessories need to be identified and reviewed along with the biopsy device

and require the same types of information as stated above.  These

accessories might include cannulas, handles/holders, introducer, spacers,

stylets, syringes, trocar, etc.  Labeling must state whether the accessory

is intended for single use and whether it is reusable or disposable.

X.   An exact identification of all materials used to fabricate the biopsy

device should be provided and a statement regarding any material

differences from the pre-amendment or substantially equivalent biopsy

device should be explicitly stated.  If the materials are identical to the

pre-amendment or substantially equivalent device and are identically

processed and sterilized, then this should be explicitly stated.  The

sponsor will need to provide biocompatibility testing data on any material

changes that have been implemented or justify why this data is not needed,

i.e. the material does not come into patient contact.  Biopsy devices are

considered to be a short-term implants, with the type of tissue it comes in

contact dependent on the indications for use.  At a minimum,

biocompatibility data should be provided for the following tests:

irritation, sensitization, cytotoxicity, acute systemic toxicity, and

implantation.  Guidance for the testing is provided in the document

entitled "Tripartite Biocompatibility Guidance for Medical Devices" and a

copy may be obtained from the Center for Devices and Radiological Health's

Division of Small Manufacturers Assistance at (800) 638-2041 or (301)

443-6597.

An exact identification of all colors (ink, dyes, markings, radiopaque

material, etc.) used to fabricate the biopsy device should be provided and

a statement regarding any colorant changes from the pre-amendment or

substantially equivalent biopsy device should be included.  If the colors

are identical to the pre-amendment or substantially equivalent device then

this should be explicitly stated.  The sponsor will need to provide

biocompatibility testing data on any colorant changes that have been

implemented; state how the markings are processed (etched, bands, in

material, etc.) and whether the color contacts skin, mucosa, etc.

XI.  The following data should be provided to demonstrate substantial

equivalence of your biopsy device with respect to functional performance.

These tests should be conducted in a manner as similar as possible to how

the biopsy device will be used in a medical procedure.  A statistically

valid number of biopsy devices should be tested to establish the

performance of each size.  A sampling of biopsy devices representative of

the product line, e.g. largest, smallest, longest, and shortest, should be

tested.  Testing should be conducted in accordance with accepted industry

standards and explicitly stated as such, or a description and analysis of

the test procedures used should be provided justifying their validity.  If

the applicant believes that functional testing is not appropriate for the

type of biopsy device to be marketed, justification should be provided.

Functional data for a cytology brush (manual or electric) must include, if

applicable, testing demonstrating electrical compatibility with other

gastroenterology-urology electrical devices; animal, bench, or other

testing that defines the appropriate specifications for motor speeds, brush

rotation, suction pressures, timing, possible electrical hazards to

patient, etc.; how the tip, if applicable, is attached to the rod and the

force needed to detach the tip, and brush/bristle attachment and strength.

Functional data for automatic biopsy devices, needles and needle sets must

include testing demonstrating needle advancement and penetration.  Accurate

needle advancement is important to ensure the sample is obtained from the

intended tissue site with minimal trauma to that site and the surrounding

tissue sites.  If claiming that the device will take multiple samples at

one site, bench testing data must be presented demonstrating that the gun

and needle assembly of the biopsy device are capable of being fired and are

able to obtain the claimed amount of samples without damage to the interior

mechanisms, and that the quality and size of the biopsy sample taken with

the later shots are equivalent for the types of tissues and/or tumors for

which the device is indicated.  If the biopsy device uses electricity,

electrical safety (isolation) must be demonstrated.

Functional data for electric biopsy forceps must include testing

demonstrating electrical compatibility with the gastroenterology-urology

electrosurgical devices recommended in the labeling, electrical safety

(isolation), and that the device can withstand the maximum currents

generated without device damage or patient and user (operator) injury.

XII. Complete information regarding biopsy devices that are sold sterile

must be provided and must include sterilization method; validation method;

packaging materials and a description of the packaging to ensure sterility

is maintained; sterility assurance level (SAL), and radiation dose or the

maximum levels of residuals of ethylene oxide, ethylene chlorohydrin, and

ethylene glycol which remain on the device, whichever applicable.  If the

device will be labeled as pyrogen free, or non-pyrogenic, provide a

description of the method used to make that determination (LAL or Rabbit

test).  If the entire biopsy device is not sold sterile, labeling must

clearly identify which parts are sterile and non-pyrogenic.  Guidance on

sterility issues is described in ODE Bluebook Memo K90-1 510(k) "Sterility

Review Guidance (2/12/90)" and a copy may be obtained from the Center for

Devices and Radiological Health's Division of Small Manufacturers

Assistance at (800) 638-2041 or (301) 443-6597.

If the biopsy device is sold and labeled nonsterile or can be reprocessed,

instructions on disassembly, cleaning, disinfection, and/or sterilization

should be provided, if applicable.  If appropriate, a statement that the

biopsy device requires high level disinfection should be provided and

compatible solutions and/or procedures for high level disinfection and/or

sterilization need to be identified.  Accessories that are disposable

should be labeled as single use.

The Association for Practitioners in Infection Control (APIC) and the

Center for Disease Control (CDC) have established definitions and

guidelines for the selection and use of disinfectants, which depend on the

site(s) of device contact.  The applicant's recommendations must be based

on the target organs stated in their labeling.

XIII. If this device is to be marketed as a kit, identify all components

and provide the certification stated below:

I certify that the following components of my kit are either (1)

legally marketed pre-amendments devices, (2) exempt from premarket

notification (consistent with the exemption criteria described in

the classification regulation(s) and the limitations of exemptions

from Section 510(k) of the act (e.g., 862.9), or (3) have been

found to be substantially equivalent through the premarket

notification process for the use(s) for which the kit is to be

intended (i.e., I am not claiming or causing a new use for the

component(s)).

I further certify that these components are not purchased in

"bulk", but are purchased in finished form, i.e., they are

packaged, labeled, etc., consistent with their pre-amendments,

exemption, or premarket notification criteria and status.

If you cannot make the above referenced certification statement

(first paragraph) for each component of your kit, you must itemize

the components without a pre-amendments, exemption, or premarket

notification status.  In this case we will continue our premarket

notification review of these components of your kit.

If you cannot make the above referenced certification statement

(second paragraph) for each component of your kit, you must

itemize these components, state whether they are pre-amendments,

exempt, or have been found substantially equivalent through the

premarket notification process, and describe how you further

process them (e.g., sterile, package/repackage, label/relabel,

etc.).

If the device kit contains components which are subject to

regulation as drugs, a substantially equivalent determination will

not apply to the drug component(s) of the device.  For information

on applicable Agency requirements for marketing the drug

component(s) in the kit, it is suggested that you contact the

Center for Drug evaluation and Research's Division of Drug

Labeling Compliance at (301) 295-8063.

https://www.fda.gov/MedicalDevices/DeviceRegulationandGuidance/GuidanceDocuments/ucm081347.htm