# Review Criteria for Assessment of Glycohemoglobin (Glycated or Glycosylated) Hemoglobin In Vitro Diagnostic Devices (Text Only)

This guidance was written prior to the February 27, 1997 implementation of FDA’s Good Guidance Practices, GGP’s. It does not create or confer rights for or on any person and does not operate to bind FDA or the public. An alternative approach may be used if such approach satisfies the requirements of the applicable statute, regulations, or both. This guidance will be updated in the next revision to include the standard elements of GGP’s.

REVIEW CRITERIA FOR ASSESSMENT OF GLYCOHEMOGLOBIN (GLYCATED or

GLYCOSYLATED) HEMOGLOBIN IN VITRO DIAGNOSTIC DEVICES.

This is a flexible document representing the current major concerns and

suggestions regarding glycohemoglobin in vitro diagnostic devices.  It is

based on 1) current basic science, 2) clinical experience and 3) previous

submissions by manufacturers to the FDA, and 4) the Safe Medical Devices

Act of 1990 (SMDA) and FDA regulations in the Code of Federal Regulations

(CFR).  As advances are made in science and medicine and changes in

implementation of Congressional legislation, these review criteria will be

re-evaluated and revised as necessary.

PURPOSE OF THE GUIDANCE DRAFT

This document is an adjunct to the 21 CFR Parts 800-1299.  While it is not

to supersede the CFR, this document provides additional guidance and

clarification on what information is necessary before the Food and Drug

Administration (FDA) can clear a device for marketing.  The FDA can make

more informed decisions based on a uniform data base.  We hope this will

lead to more reliable, reproducible and standardized commercial tests.

DEFINITION

This document discusses all generic type devices intended for use in

clinical laboratories as in vitro diagnostic tests for the quantitative

determination of the fraction (expressed as per cent) of glycohemoglobin

(glycated or glycosylated) hemoglobin (GHb%) in blood samples.1

PRODUCT CODE: LCP

REGULATION NUMBER: 21 CFR   864.7470

CLASSIFICATION. Class II

PANEL:  Hematology (81)

REVIEW REQUIRED:  Premarket Notification (510(k).

1.0 CLINICAL INDICATION/SIGNIFICANCE/INTENDED USE OF GLYCOHEMOGLOBIN

TESTS.

1. Clinical Indication/Significance/Intended Use

Glycohemoglobin (GHb) is being used with increasing frequency to monitor

long-term blood glucose control and compliance in patients with diabetes

mellitus.  The GHb test provides an index of the mean concentration of

blood glucose during the preceding two to three months.  It complements

more traditional measures of glucose control, such as glucose quantitation

in urine and blood.2

Diabetes mellitus is a metabolic disorder characterized by impairment of

carbohydrate metabolism.  Two major types have been identified.  In type I

insulin dependent diabetes) the pancreatic B cells fail to produce insulin,

the protein hormone necessary for glucose oxidation to take place within

cells of the body.  Type II or Non-Insulin-Dependent-Diabetes Mellitus

(NIDDM) patients are of two subtypes: nonobese and obese. Nonobese Type II

diabetics have impaired insulin production.  Type II obese NIDDM patients

manifest ineffective insulin stimulation of target cells which is believed

to be a post-receptor defect in target tissues. The different types of

diabetics present with different signs and symp-toms,  but all have

characteristic fasting hyperglycemia or elevation of glucose at 1 to 2

hours.  The goal of treatment of diabetes mellitus has been to control

blood glucose levels to as near normal as possible in the hope that this

will impede, if not prevent, progression of vascular disease, e.g., stroke,

blindness, renal failure, heart attack, circulatory insufficiency to lower

extremities, etc.  This may be accomplished through diet, weight loss, oral

hypoglycemic drugs or by regular injections of insulin, depending on the

type of diabetes and individual patient differences.3,4

The term "glycohemoglobin" (GHb) refers to a series of minor hemoglobin

components that are stable adducts formed by hemoglobin with various

sugars.  The reaction between glucose and hemoglobin is an example of a

nonenzymatic condensation of glucose with the free amino groups on the

globin (protein) component of hemoglobin1.  This process is slow,

continuous and irreversible.  The human erythrocyte is freely permeable to

glucose.  Within each erythrocyte, GHb is formed from hemoglobin at a rate

dependent on the ambient concentration of glucose.5  The higher the

prevailing ambient levels of blood glucose, the higher will be the levels

of GHb.  Per cent GHb is increased within the erythrocytes of patients with

diabetes mellitus.6

Tests for glycohemoglobin are unreliable for the diagnosis of diabetes

mellitus.  Such use yields too many false negative and false positive

results.2  It may be useful, however, to screen for those at greatest risk

for diabetes complications.  It is not yet clear if patients diagnosed as

having diabetes based on oral glucose tolerance test results with

persistent normal per cent glycohemoglobin GHb%) values develop typical

diabetic complications.  A claim for GHb tests to diagnose diabetes

mellitus then would require premarket approval.

False results could suggest incorrect long term blood glucose control. When

used with the more immediate testing of glucose in blood and urine,

however, false GHb% test results can be evaluated before disastrous

measures are taken, because there could be short term low or high glucose

levels in blood and/or urine with high or low GHb% levels (long term

effect).

Specimen types:

List all specimen types/matrix(ices) claimed by the manufacturer for use in

the test.  Matrix is defined as the milieu containing the analyte in the

patient sample submitted for analysis (hereinafter "specimen type" includes

considerations for matrix or milieu)15.  It can be whole blood, packed

erythrocytes and/or lysed washed packed red cells, etc.

2. DEVICE DESCRIPTION:

Discuss the principles of the device methodology and whether it is

well-established or new and unproven.

3. NONCLINICAL LABORATORY STUDIES: SPECIFIC PERFORMANCE CHARACTERISTICS

FDA requests different types and amounts of data and statistical analyses

in applications to market in-vitro diagnostic devices.  The amount and type

of data requested depends on: 1) the test analyte, 2)  the intended use

(which determines whether the application is a 510(k), an original

Premarket Approval application (PMA), 3) whether the test is quantitative

or qualitative, 4) whether the data design is independent or paired, and 5)

certain claims made by the manufacturer.

The performance of the device can be established by comparison of the

device to a predicate device (any legally marketed device).  The National

Committee for Clinical Laboratory Standards (NCCLS) is an example of a

source for reference methods.  Prove all claims for substantial equivalence

and specific parameters for using the device.  Include data to support use

of the test with all claimed specimen types/matrices.

A. Analytical/Laboratory/In Vitro Studies.

1. Statistical

Describe statistical methods, assumptions used, statistical analyses and

corresponding computer outputs and references well enough so a

knowledgeable reader with access to the original data can verify the

reported results.

Provide data and statistical analyses determined with the device to support

performance parameters specific to and important for operating the device,

e.g., reproducibility.

Discuss the statistical methods for the type of data submitted, e.g.,

quantitative continuous data, qualitative discrete data, etc.  The

distribution of data (normal vs. Non-normal type of data (paired vs.

independent).

Use references for study design and statistical methods that are from

standard texts and/or refereed biomedical journals.

Test data:

2. Performance Characteristics

Establish the linearity of the test in relation to the hemoglobin

concentration of the sample.  Using varying concentrations of hemoglobin,

document the range of hemoglobin in which accurate GHb% results are

obtained.

In the package insert state the hemoglobin levels which may cause the test

to become over- or under-loaded and give aberrant results.  Give

instructions in the package insert for the action to be taken when samples

are above or below linearity ranges.

3. Specificity/Cross-Reactivity/Interference Studies

If the manufacturer makes a claim for restricted assay specificity, e.g.,

it is claimed that an immunologically-based test, electrophoresis or

cation exchange methodology detects HbA1c rather than HbA1 only, the

specificity of that test for the claimed restricted species should be

demonstrated.

For Tests Employing Cation Exchange, HPLC and Electrophoresis

Methodologies, present chromatograms from gas-liquid and liquid

chromatography or electrophoregrams from electrophoretic techniques so

that users can see the efficiency of the separation of the claimed

glycohemoglobin peak from other peaks and observe the resolution from

interfering substances in the matrix.  Express migration by Rf (or

retention times for columns).  State temperature conditions.  Describe in

detail the solvents or carriers and their order of use.17  Include similar

types of information for electrophoresis as for chromatography as

applicable.17

Present chromatograms or electrophoregrams of samples containing

Hemoglobin F to determine whether of not there is interference.

Present chromatograms or electerophoregrams of samples from patients with

various hemoglobinopathies (such as HbS, HbG, HbH, Hb Wayne, HbC,

thalassemia, etc., to determine whether or not there is interference.

Describe the exact source of columns and packing materials, and their

handling.  Give the solvent sources and solvent compositions, and describe

the history of experience with each column.

Data should be provided to demonstrate that the specimen does not exhibit

test interference if it is hemolyzed, lipemic (affinity chromatography

only) or bilirubinemic at specified levels24.  Any of these conditions

found to interfere with the test should be declared as interfering

conditions in the package insert.

It is already documented that electrophoresis tests are unaffected by

hypertriglyceridemia25 and that tests employing cation exchange methodology

are affected by severe lipemia26 and bilirubinemia27.  HPLC is also

affected

by bilirubinemia27.  If these limitations are declared and referenced in

the respective package insert, the respective data does not have to be

provided.

Labile GHb is an acutely generated, non-enzymatic, reversibly linked

glucose intermediary product present in blood after a heavy meal which can

cause a falsely elevated GHb% test.

Submit data to support a claim that the presence of labile GHb does not

cause a statistically significant elevation of the true GHb% value.

Several test protocols are acceptable for this purpose.

Show that normal and diabetic specimens incubated for 3 to 4 hours at 37oC

with at least 55 Mm/L or 1400mg/Dl of glucose28,29 give the same GHb%

values

as the original untreated specimen.

If the test cannot meet this stringent, non-physiological test, it will be

acceptable to show that for 21 samples in the high GHb range there is not

a clinical statistically significant difference between GHb% values

obtained with the new test vs. GHb% values obtained using an accepted

(published) labile removal method analyzed by paired student's t test, and

that the "labile-removed" value is lower than that with no labile removal,

i.e., labile GHb was present in the sample.  For example, take 21 samples

with HbA1c values between 10 and 15%.  Divide each sample into 3 aliquots:

a) aliquot #1 - no treatment, b) aliquot #2 remove labile by the method

being tested, c) aliquot #3 - remove labile by incubation of washed

erythrocytes in saline at 37oC for 5 hours or by another accepted

(published) method.  Run aliquots #1, #2 and #3 in the same assay.

Aliquot #2 and #3 should give the same results and be lower than #1 (or

equal to #1 if the assay itself is not affected by labile GHb).

If an interference is not examined, then add a statement to the

Limitations Section of the package insert that the device has not been

tested for cross-reactivity or interference of one or more substances.

Analysis of Variance of Reproducibility17,19,20,22,23

The National Committee for Clinical Laboratory Standards (NCCLS)

recommends23 an analysis of variance experiment testing two clinically

significant levels near medical decision limits (subnormal, normal, or

elevated) of an analyte, in this case normal and elevated.  Use controls

simulating patient samples or actual patient specimens 2 times in the same

run and in two different runs each day for 20 days.  This permits separate

estimation of between-day, between-run and within-day standard deviations

(SD), as well as within-run and total SDs.  Acceptable alternatives that

include only one run per day are also discussed in the cited document.23

The three important assumptions (homogeneity of error variances,

additivity, and normality) of analysis of variance should be used to

demonstrate the validity of the above results.

Calculate total, between- and within-day, and between- and within-run

means and coefficients of variation of imprecision for each set of values.

5. Comparison Studies.

Compare the device to a FDA-cleared device.  In addition the device may be

compared to a reference method.  GHb A1c by HPLC or an affinity

chromatography method is recommended for use as a reference method14.  HPLC

products are commercially available.  HPLC shows excellent assay precision

and permits rapid separation of HbA1c from the other minor components.3

The A1c peak suffers less variability from sample storage does the HbA1

peak.4 Affinity chromatography has far fewer interferences including sample

degradation problems than do any of the ion-exchange or electrophoresis

methods.  The reference method chosen should be used for all comparisons

throughout the study.

Compare results obtained using packed red cells (if claimed) and/or whole

blood samples free from interfering substances from 40 to 100 persons

covering the whole assay range (from normal to clinically relevant high

levels of GHb%) to results obtained with another test already on the market

or to a reference method.17,18

Analyze the data using linear regression methods18. (The X axis is the

independent variable or comparison test.  The Y axis is the dependent

variable or new test.)19,20  Linear-regression analysis is often most

useful for estimating the differences or errors between two analytical

methods, because the errors can be calculated at any medically important

concentration within the range studied; furthermore, the slope and

intercept may give some indication of the type of systematic error, which

may aid in reducing the analytical errors.  Because the reliability of the

estimates of slope and intercept can be affected by nonlinearity in the

data set, outliers, a narrow range of data, and variability of the

comparison method, samples preferably should cover the complete range of

concentrations that might be encountered.21  The slope, intercept, and

their estimated standard errors, correlation coefficient, the standard

error of the estimate, the assay range, and nature and size of samples

tested should be reported in the Performance Characteristics section of the

package insert.

B. Clinical Data/Reference Ranges.

The device results may correlate well using linear regression (slope close

to 1.0 and intercept close to zero)17,18 with a method that has a published

reference range for healthy individuals.  If the device results correlate

well, 40-60 subjects are sufficient to confirm agreement.30,31

If the device results do not correlate well, establish a reference range

with samples from 120 to 200 normal persons characterized by age, sex,

geographic location, any symptoms of disease and any other factors that

would influence the values obtained, e.g., pregnancy.31

It is suggested that the statistics used to characterize the population and

the confidence limits used be stated in the package insert.  The

populations studied should be characterized according to age and disease

status.  The number of persons tested should be stated.  Investigate all

sample type(s) claimed in the intended use unless other data proves that

there is no difference between them.  A range of analyte values for samples

from specified patient groups, e.g., diabetics, may also be provided.

4. LABELING CONSIDERATIONS

Do not make unproven claims for clinical significance in the package

insert.

The following are additional details for labeling.

A. The Intended Use Statement [ 809.10(b)(2)]

Whether it is for use in clinical laboratories, doctors' offices, or over-

the-counter (OTC). (The Limitations section should include any specific

training required for test performance or use.)

Whether it is for screening, monitoring, confirmation or exclusion and/or

to aid in the diagnosis as an adjunct to other procedures.

Clinical significance, if it can be stated in a few words.  (If the

clinical significance statement is lengthy or complicated, create a

separate heading entitled "Clinical Significance".)

A typical intended use statements:

"ABC's *** test is a laboratory test intended for the quantitative

determination of percent glycated hemoglobin in whole blood by

[methodology] using the ABC automated system (if applicable) to monitor

long term blood glucose control in individuals with diabetes mellitus."

Conditions for Use

Describe any special applications of the device or specific

contraindications or indications for use not addressed in the Intended Use

Statement, e.g., "Despite serious consideration the measurement of

glycohemoglobin has not been shown to be reliable in the diagnosis of

diabetes mellitus."2,5  For example, one study shows that even though a

cutoff of three standard deviations above the mean of a "normal" population

has a specificity of 99% for diabetes, the sensitivity is only 48%49.

B. Summary and Explanation of the Test [ 809.10(b)(3)]

It is suggested that this section should discuss the following merits and

limitations of this test:

1. For Affinity Chromatographic Methods Only;

This method detects all glycohemoglobins, not just HbA132,

This method is not affected by the presence of abnormal hemoglobins.

This method is unaffected by carbamlyated hemoglobins in uremic patients.

In uremic patients there is condensation of urea-derived cyanate with the

N-terminal amino groups on the beta chains of HbA.  This urea bound

hemoglobin elutes with HbA1 on cation-exchange columns producing falsely

elevated results33,34.

2. Labile Glycohemoglobin (All methods, if data so demonstrates.)

This method is unaffected by the presence of "labile" glycosylated

hemoglobin.

C. Specimen collection and preparation for analysis [ 809.10(b)(7)]

Data or literature references should be provided in the 510(k) to back all

statements made.

Include a description of:

The type of specimen to be collected, e.g., whole blood, packed cells,

washed packed cells, and acceptable anticoagulants.

Acceptable anticoagulants or other additives, preservatives, etc., to

maintain specimen.

Collection Precautions:

Special conditions for patient preparation, e.g., fasting, timing of

collection, intervals of collection, etc.

For routine clinical use, testing every 3 to 4 months is generally

sufficient.  In certain clinical situations such as diabetic pregnancy, or

after a major change in therapy, it may be useful to obtain GHb% values in

2 to 4 week intervals5.

Electrophoresis Methodology

A statement that hypertriglyceridemia does not interfere electrophoresis

methodology25.

For Tests Employing Cation Exchange Methodology, severely lipemic samples

may show elevated results26. If a sample appears lipemic, it is recommended

that washed packed red cells be used as the sample.

For Tests Employing Cation-Exchange or HPLC Test Methodology, it has been

reported that samples with elevated levels of bilirubin may exhibit

falsely elevated levels of glycohemoglobin27.

State in the package insert in range of oC (low to high), collection,

transport, handling and storage conditions to maintain stability of the

specimen.  Do not use terms such as "room temperature" without

qualification.  Provide data or appropriate literature references in the

510(k) to back any claims made.

Due to instability of the samples for Cation-Exchange and Electrophoresis

methodologies consistent storage conditions and time intervals between

samples should be recommended.  For example, all samples should be

prepared into hemolysates on the same day of collection and tested on the

morning of the fourth day after collection.

A statement that each laboratory must develop and evaluate sample handling

procedures based on their specific situation should be made.  The

laboratory should be responsible for educating clinicians concerning the

large analytical errors that can be introduced by inappropriate sample

handling.

[Storage conditions may contribute significantly to both cation-exchange

and electrophoresis assay imprecision6

The temperature variation for some cation exchange minicolumns may be 1%

HbA1 per 1oC46.  This temperature problem must be addressed in any one of

several ways, for example:

However, reports indicate that a temperature vs. GHb concentration

conversion chart47 may not provide accurate results6.  Such charts do not

compensate for temperature fluctuations during assay11.

Use three level calibrators to compensate for temperature variations and

recommendations for methods of strict temperature control.

Various substances other than sugars can form adducts with hemoglobin,

thereby altering its charge characteristics.  Falsely elevated results can

occur if these adducts comigrate with GHb.  Examples include individuals

with opiate addiction44, lead poisoning, uremia, and alcoholism,5 as well

as those receiving large doses of aspirin (acetylated hemoglobin).45

Clinically the most important of these interfering adducts occurs in

uremia.5  This increase correlates with the BUN (blood urea nitrogen) and

appears to result at least in part from the carbamylation of hemoglobin by

urea-derived cyanate (carbamylated hemoglobin).33

D. Quality Control [ 809.10(b)(8)(vi)]

Describe specimens or commercially available products that should be used

for positive and negative control including recommended levels of analyte,

if materials are not provided in the kit.

Recommendations for frequency and placement of quality control samples

within run and from run to run.  Directions for interpretation of the

results of quality control samples (satisfactory limits of performance).

Conclude with a statement similar to the following: "If controls do not

behave as stated (above), test results are invalid."

E. Limitations of the Procedure [ 809.10(b)(10)]

The following are examples of the types of limitation statements for GHb%

that are suggested for inclusion:

1. Limitation applying to all GHb% test methods

This test is unreliable for the diagnosis of diabetes mellitus2,5.  There

are too many false positive and/or false negative results, depending upon

where the test cutoff is set.  For example, one study shows that even

though a cutoff of three standard deviations above the mean of a "normal"

population has a specificity of 99% for diabetes, the sensitivity is only

48%49.

This test is not useful in judging day-to-day glucose control, and should

not be used to replace daily home testing of urine and blood glucose5.

As with any other laboratory procedure, a large discrepancy between

clinical impression and test results usually warrants investigation.  Some

of the following test limitations should be considered.

Any cause of shortened red cell survival will reduce exposure of red cells

to glucose with a consequent decrease in %Ghb values, e.g., hemolytic

anemia or other hemolytic diseases, pregnancy, recent significant blood

loss, etc.  Per cent Ghb results are not reliable in any patients with

chronic blood loss and consequent variable erythrocyte

lifespan6,11,36,37,38.

2. All tests except those with data demonstrating lack of interference by

labile Ghb. [If the manufacturer claims that the methodology of the test

removes or is unaffected by labile glycohemoglobin, data should be

presented in the 510(k) to back the claims.]  Falsely elevated %Ghb test

results can be caused by "labile Ghb", an acutely generated, reversible,

non-enzymatically linked glucose intermediary product present after a heavy

meal.  The higher the prevailing ambient levels of blood glucose, the

higher will be the potential for falsely elevated results caused by labile

Ghb.  The package insert should advise users that labile Ghb may be removed

by any one of the following methods, depending on the test methodology

employed:

1) Incubation of a twenty-fold sample dilution overnight in isotonic

saline at room temperature39.

2) Dialysis overnight at 4oC vs. two changes of volumes of buffer

containing 4.59 g  NaH2PO4.H2O, 1.18 g  Na2HPO4 and 0.65 g KCN per liter

(pH

7.0)40.

3) Incubation at 37oC for 5 hours in 0.9% saline41.

4) Incubation for 30 minutes at 38oC in 30 mM semicarbazide and 12 mM

aniline, pH 5.  Note: These chemicals are toxic28.

5) Lysis of erythrocytes for 15 minutes in 50 volumes of 0.05M potassium

biphthalate lysing buffer, Ph 5 at 37oC.  Care must be taken to avoid

denaturation of hemoglobin29,42.

3. For all cation exchange (including HPLC), and electrophoresis

methodologies

Note that elevated levels of Hemoglobin F (HbF) usually found in infants

and some pregnant women result in falsely elevated results because HbF

comigrates with HbA1c43.

Patients with various hemoglobinopathies (such as HbS, HbG, HbH, Hb Wayne,

HbC, thalassemia, etc.) may have incorrect results depending on charge

characteristics of these variants6.

4. For assays quantifying HbA1 only, note that in most of the following

situations, including uremia, the HbA1a and b fraction is more affected

than the HbA1c fraction. Thus assays that quantify HbA1c specifically show

only slight alterations (rarely greater than 1 GHb%) from these

interferences.5

F.Interpretation of Results/Expected Results [ 809.10(b)(11)]

Actions to be taken if a samples' test result is above or below linearity

ranges according to hemoglobin concentration, e.g., dilution procedures.

Ranges for healthy persons

Ranges for defined disease groups.

Reference ranges from the scientific literature

Alternatively, reference range studies may also be reported from

literature references, especially if the new test was used by the authors,

or if the performance characteristics demonstrate results equivalent to

those obtained with a test used by the author, e.g., regression equation

slope = 1.0, r = approximately 0.95.  Otherwise all expected values should

be determined with the submitted product.

In normal healthy individuals,  GHb comprises approximately 4 to 8% of the

total hemoglobin.  This level may be doubled in the diabetic patient.

Because glucose binding to hemoglobin occurs slowly and depends on the

circulating level of blood glucose, the GHb% level represents a time-

averaged blood glucose level.  There is a time lag of approximately two to

four weeks before GHb% reflects changes in the blood glucose level.

Known insulin-dependent diabetic patients usually have elevated GHb%

levels48. Percent GHb quantitations on these patients may be in the 9-17%

range depending on the degree of hyperglycemia.  Diabetic patients in good

control may have GHb% values in the normal range.  To date, there is no

specific GHb% value that is accepted as indicating "good" or "poor"

control.  When using GHb% to monitor the diabetic patient, results must be

interpreted individually; that is, the patient should be monitored against

him- or herself and values compared to the normal range for that

particular method.

For Electrophoresis and High Performance Liquid Chromatography Test

Methodology.  Provide Graphics

Provide examples of electrophoregrams from common hemoglobinopathies,

e.g., HbS and C.  Provide a detailed discussion with examples of how to

calculate GHb% in the presence of genetically abnormal hemoglobins.

G. Performance Characteristics [  809.10(b)(12)]

Provide the statistical results of the comparison study.  Scattergrams are

recommended but optional.

Report the mean, standard deviation and/or per cent coefficient of

variation for the various levels of %GHb in the within-run, within-day,

between-day and total assay precision studies.  Mean values should be

arranged from lowest to highest (or vice-versa) to illustrate recognizable

trends.

Summarize studies performed to establish assay linearity depending on

hemoglobin concentration, and state conclusions.

5. BIBLIOGRAPHY

1. Fuentes-Arderiu X. "Glycohemoglobin," not "glycated hemoglobin" or

"glycosylated hemoglobin"  Clin Chem 1990;36:1254.

2. Lester E.  The clinical value of glycated hemoglobin and glycated

plasma proteins.  Ann Clin Biochem 1989;26:213-9.

3. Karam JH.  Diabetes mellitus, hypoglycemia, & lipoprotein disorders.

In Schroeder SA, Krupp MA, Tierney LM Jr, McPhee SJ, editors.  Current

medical diagnosis and treatment 1991.  Norwalk and San Mateo: Appleton &

Lange, 1991:852-893.

4. The National Diabetes Data Group. Diabetes 1979;28:1039

5. Goldstein DE, Little RR, Wiedmeyer HM, England JD, and McKenzie EM.

Glycated hemoglobin: methodologies and clinical applications.  Clin Chem

1986;32:B64-B70.

6. Goldstein DE, Wiedmeyer HM, England JD, Little RR, Parker KM. Recent

advances in glycosylated hemoglobin measurements. CRC Crit Rev Clin Lab

Sci 1984;21:187-228.

7. Larsen ML, Hþrder M, Mogensen EF. Effect of long-term monitoring of

glycosylated hemoglobin levels in insulin-dependent diabetes mellitus.

New Engl J Med 1990;323:1021-1025.

8. The DCCT Research Group. Diabetes control and complications trial

(DCCT); Update.  Diabetes Care 1990;13:427-33.

9. Kunkel HG and Wallenuis G. New hemoglobins in normal adult blood.

Science 1955;122:288.

10. Bookchin RM, Gallop PM. Struycture of hemoglobin A1c: nature of the N-

terminal beta chain blocking group.  Biochem Biophys Res Commun

1968;32:86-93

11. Peacock I. Glycosylated haemoglobin: measurement and clinical use.  J

Clin Pathol 1984;37:841-51.

12. Mallia AK, Hermanson GT, Krohn RI, Fujimoto EK, Smith PK. Anal Lett

1981;14:649-61.

13. The DCCT Research Group.  Feasibility of centralized measurements of

glycated hemoglobin in the diabetes control and complications trial: a

multicenter study. Clin Chem 1987;33:2267-71.

14. Little RR, England JD, Wiedmeyer HM, McKenzie EM, Mitra R, Erhart PM,

Durham JB, and Goldstein DE. Interlaboratory standardization of glycated

hemoglobin determinations.  Clin Chem 1986;32:358-60.

15. Nomenclature and definitions for use in the national reference system

for the clinical laboratory. 5(21):561.  Order code NRSCL8-P, ISBN 0273-

3099

16. International Committee of Medical Journal Editors. Special report.

Uniform requirements for manuscripts submitted to biomedical journals.

NEJM 1991;324:424-428.

17. Information for authors. Clin Chem 1991;37:1-3.

18. National Committee for Clinical Laboratory standards. User Comparison

of quantitative clinical laboratory methods using patient samples,

proposed guideline. 1985;6(1). Order code EP9-P.

19. Peters T, Westgard JO. Evaluation of methods, Chapter 7 in: Tietz NW,

editor. Fundamentals of clinical chemistry, Third Edition, Philadelphia:

Saunders. 1987: 225-37.

20. Westgard JO, de Vos DJ, Hunt MR, Quam EF, Carey RN, Garber CC.  Method

evaluation.  American Society of Medical Technology, Bellaire, TX, 1978.

21. Information for authors. Clin Chem 1990;36:1-4.

22. Vadlamudi SK, Stewart WD, Fugate KJ, and Tsakeris TM. Performance

characteristics for an immunoassay.  Scand J Clin Lab Invest 1991;51:134-

138.

23. National Committee for Clinical Laboratory Standards. Evaluation of

precision performance of clinical chemistry devices - second edition;

tentative guidelines. 1991:1-56. Order Code EP5-T2.

24. National Committee for Clinical Laboratory Standards.  Interference

Testing in clinical chemistry; proposed guideline. 1986  Order code EP7-P

25. Aleyassine H, Gardiner RJ, Blankstein LA, Dempsey ME. Agar gel

electrophoretic determination of glycosylated hemoglobin: effect of

variant hemoglobins, hyperlipidemia, and temperature. Clin Chem

1981;27:472.

26. Dix D, Cohen P, Kingsley S, Lea MJ, Senkbeil J, Sexton K. Interference

by lactesence in glycohemoglobin analysis.  Clin Chem 1979;25: 494-5.

27. Simon M, Eissler J. Critical factors in the chromatographic

measurement of glycohemoglobin (HbA1).  Diabetes 1980;29: 467-74.

28. Nathan DM, Avezzano ES, and Palmer JL. A rapid chemical means for

removing labile glycohemoglobin.  Diabetes 1981;30:700-1.

29. Biss‚ E, Berger W, Flckiger R. Quantitation of glycosylated

hemoglobin, elimination of labile glycohemoglobin during sample hemolysis

at pH 5.  Diabetes 1982;31:630-3.

30. Ash, KO. Reference intervals (normal ranges): a challenge to

laboratorians. Am J Med Tech 1980;46:504-11.

31. National Committee for Clinical Laboratory Standards.  How to define,

determine, and utilize reference intervals in the clinical laboratory;

proposed guideline. Villanova, PA. 1991. Order code C28-P.

32. Klenk DC, Hermanson GT, Krohn RI, Fujimoto EK, Mallia AK, Smith PK,

England JD, Wiedmeyer HM, Little RR, Goldstein DE. Determination of

glycosylated hemoglobin by affinity chromatography: comparison with

colorimetric and ion-exchange methods, and effects of common

interferences.  Clin Chem 1982;28: 2088-94.

33. Fluckiger R, Harmon W, Meier W, Loo S, Gabbay KH.  Hemoglobin

carbamylation in uremia. N Eng J Med 1981;304: 823-7.

34. Bruns DE, Lobo PI, Savory J, Wills MR. Specific affinity-

chromatographic measurement of glycated hemoglobins in uremic patients.

Clin Chem 1984;30: 569-71.

35. National Committee for Clinical Laboratory Standards.  Internal

Quality Control Testing:  Principles and Definitions; approved guideline.

Villanova, PA. 1991. Order code C24-A:4

36. Bunn HF, Gabbay KH, Gallop PM.  The glycosylation of hemoglobin:

relevance to diabetes mellitus.  Science 1987;200:21-7.

37. Horton BF, Huisman THJ.  Studies on the heterogeneity of haemoglobin,

VII; minor haemoglobin components in haematological disease.  Brit J

Haemat 1965;11:296-304.

38. Lind T, Cheyne GA.  Effect of normal pregnancy upon the glycosylated

hemoglobin.  Br J Obstet Gyn 1979;86:210-3.

39. Shenouda FS, Cockram CS, Baron MD, Tsatsoulis A, Wen Han L, Sonksen

PH. Importance of short-term changes in glycosylated haemoglobin. Br Med

J 1982;284:1084-5.

40. Trevilli LA, Ranney HM, Lai H.  Hemoglobin components in patients with

diabetes mellitus.  New Engl J Med 1971;284:353-7.

41. Goldstein DE, Peth SB, England JD, Hess RL, Da Costa J. Effects of

acute changes in blood glucose on HbA1c Diabetes 1980;29:623-8.

42. Bannon P. Effect of pH on the elimination of the labile fraction of

glycosylated hemoglobin.  Clin Chem 1982;28:2183.

43. Menard L, Dempsey ME, Blankstein LA, Aleyassine H, Wacks M, Soeldner

JS. Quantitative determination of glycosylated hemoglobin a1 by agar gel

electrophoresis. Clin Chem 1980;26:1598.

44. Coriello A, Giugliano D, Dello Russo P, Sgambato S, D'Onotrio F.

Increased glycosylated hemoglobin A1 in opiate addicts.  Evidence for

hyperglycemic effect of morphine.  Diabetologia 1962;22:379.

45. Nathan DM, Francis TB, Palmer JL.  Effect of aspirin on determinations

of glycosylated hemoglobin.  Clin Chem 1983;29:466-9.

46. Schellekens APM, Sanders GTB, Thornton W, von Groenestein T. Sources

of variation in the column-chromatographic determination of

glycohemoglobin (HbA1). Clin Chem 1981;27: 94-9.

47. Hankins WD, Holladay L.  A temperature conversion nomogram for

glycosylated hemoglobin analysis. Clin Chem Acta 1980;104: 251.

48. Rahbar S. Biochem Biophys Res Commun 1969;36:838-43.

49. Forrest RD, Jackson CA, Yudkin JS. The glycohaemoglobin assay as a

screening test for diabetes mellitus: the Islington Diabetes Survey.

Diabetic Med. 1987;4: 254-9.

HFK-440 NChace/chron  2/24/91 Version 9/27/91

This document was still considered current as of July 1997.
It will be reviewed again in July 1998.