# Assessing the Safety and Effectiveness of Home-Use In Vitro Diagnostic Devices (IVDs): Draft Points to Consider Regarding Labeling and Premarket Submissions (Text Only)

This guidance was written prior to the February 27, 1997 implementation of FDA’s Good Guidance Practices, GGP’s. It does not create or confer rights for or on any person and does not operate to bind FDA or the public. An alternative approach may be used if such approach satisfies the requirements of the applicable statute, regulations, or both. This guidance will be updated in the next revision to include the standard elements of GGP’s.

Assessing the Safety and Effectiveness of Home-Use

In Vitro Diagnostic Devices (IVDs):

Draft Points to Consider

Regarding Labeling and Premarket Submissions

Center for Devices and Radiological Health

October 1988

I.   Purpose

The purpose of this document is to assist prospective

manufacturers, producers, and marketers (hereinafter called

manufacturers) of home-use in vitro diagnostic devices

(IVDs) in complying with existing labeling regulations,

e.g., 21 CFR 801.5, 809.3(a) and 809.10, and premarket

clearance requirements, e.g., Sections 510(k) and 515 of the

Federal Food, Drug, and Cosmetic Act (the Act), and

regulations including 21 CFR Parts 807 and 814.  This

document addresses home-use testing kits and home-use

mail-in specimen collection kits.

II.  Background

In vitro diagnostic devices (IVDs) as defined in 21 CFR

809.3(a) are those reagents, instruments, and systems

intended for use in the diagnosis of disease or other

conditions, including a determination of the state of

health, in order to cure, mitigate, treat, or prevent

disease or its sequelae.  These products are intended for

use in the collection, preparation, and examination of

specimens taken from the human body.

Traditionally, IVDs have been used primarily by hospitals,

clinical laboratories and physicians' offices.  In recent

years, however, there has been increasing interest in

home-use IVDs, where no physician is involved to interpret

the test results.

Because of this growing interest, FDA's Center for Devices

and Radiological Health (CDRH) anticipates receiving an

increasing number of premarket product submissions, i.e.,

premarket notifications or premarket approval applications,

for these devices.  Consequently, CDRH believes that uniform

evaluation criteria for home-use IVDs are needed to better

ensure that these devices are regulated in a consistent

fashion and that consumers are provided with reliable,

useful, and adequately labeled products.  To this end, in

the spring of 1985, CDRH solicited views from various

industry, consumer, and health professional organizations to

help identify the issues to be resolved by CDRH in

developing evaluation criteria for home-use IVDs.  In

addition, on September 9, 1985, CDRH convened a public

meeting of the chairpersons, the consumer representatives,

and the industry representatives of the Immunology Devices

Panel, the Microbiology Devices Panel, the Hematology and

Pathology Devices Panel, and the Clinical Chemistry and

Clinical Toxicology Devices Panel. The purpose of this

meeting was to solicit the views of these individuals and

other interested persons regarding general issues of the

safety and effectiveness of home-use IVDs.

The present document reflects CDRH's views on key points to

be considered in establishing the safe and effective use of

home-use IVDs.  These views are significantly based on the

input CDRH has received from the aforementioned outreach

activities.

III. Factors CDRH will Consider in Evaluating the Safety and

Effectiveness of Home-use IVDs

A.   Performance Considerations

A key factor CDRH must consider in assessing the safety

and effectiveness of home-use IVDs is test performance.

Because the great majority of proposed home-use IVDs

are likely to be derived from professional-use IVDs, it

is expected that there will be a clear and established

relationship between the analyte to be tested using the

home-use IVD and a specific medical disorder or

condition.  Because home-use IVDs function essentially

the same as professional-use IVDs, their performance

characteristics can be defined in terms of traditional

performance parameters such as sensitivity,

specificity, accuracy, and reproducibility of test

results. Nevertheless, the performance of the device in

the hands of skilled users may not reflect the device's

performance in the hands of lay users.  Consequently,

manufacturers of home-use IVDs should focus their

attention on the following performance considerations:

1.   For a given intended use, the ability of the

home-use IVD to detect and/or measure a particular

analyte of interest (analytical performance)

should be comparable to the performance of IVDs

intended for the same purpose that are used by

professionals in clinical settings;

2.   Home-use IVDs should be designed with a view to

ensuring that the device's performance will not be

appreciably affected by anticipated variation in

user technique; and,

3.   Home-use IVDs should include a simple method by

which the user can reasonably verify that the

product's performance meets its design

specifications at the time of use.  Ideally, a

user quality control test should be provided or

"built into" each home-use IVD (see 21 CFR

809.10(a)(6).)  Omission of such a test or other

reasonable means of verifying product performance

should be justified based on scientific merit.

B.   Risk/Benefit Considerations

Home-use IVDs are different from clinical laboratory

devices in four important ways:  (1) the person

performing the test or reviewing the results returned

from a clinical laboratory likely lacks the necessary

medical training to evaluate other related medical

information, such as, pertinent personal or family

medical history, other analyte levels, and the results

of a complete physical examination; (2) the person

performing the test may lack the technical training of

laboratory technologists who traditionally have

performed laboratory tests and may not interpret

directions carefully; (3) the person performing the

test may or may not carry out necessary follow-up

actions on the basis of test results; and (4) the

person collecting a sample of a body specimen and

mailing it to a clinical laboratory generally lacks the

technical training of persons who traditionally have

collected or supervised the collection of body

specimens, and supervised the transport of the

specimens to the test laboratory.  Thus the person may

not ensure that the specimen that arrives at the

clinical laboratory is in the same condition as it was

when it was collected (e.g. ensuring that the analyte

to be measured does not degrade because of passage of a

long period of time from specimen collection until

analysis, or ensuring that a specimen is not exposed to

extremes of hot or cold temperatures or other

environmental hazards that may affect the test

results).

Hence, in evaluating home-use IVDs, CDRH must consider

probable health benefits to be gained in the use of

such devices compared to the probable risks associated

with their use.  In this context, consideration of

benefit and risk is inherent in the evaluation of

safety and effectiveness.

When the use of an in-home IVD can be demonstrated to

be substantially equivalent, in terms of safety and

effectiveness, to a clinical laboratory device when

used by a health professional, and the device meets all

labeling and other regulatory and statutory

requirements, it will generally satisfy requirements

for market clearance.  When unique issues of safety and

effectiveness arise and substantial equivalence cannot

be demonstrated, CDRH will nonetheless evaluate benefit

and risk factors in making determinations with respect

to the marketability of in-home IVDs.  For purpose of

assessing these benefit and risk factors CDRH will

consider the following questions and encourage

applicants to provide relevant information in their

submissions.

Benefits

1.   What is the clinical benefit of the test to the

patient or society (public health) in terms of

screening, diagnosing, or monitoring a particular

disease, condition, or risk factor?

2.   What are the benefits to the patient or society of

having the test available for home-use as opposed

to having the test performed only by health care

professionals?

Risks

1.   What is the impact on the user or to society of a

false-positive or false-negative test result e.g.,

in terms of user follow-up or adverse medical

conditions?

2.   What are the risks to the user or society in terms

of delay in obtaining a professional examination

if a proposed home-use IVD that is intended for

use on symptomatic subjects gives a false or

equivocal result?

CDRH recognizes that for evaluation purposes, responses

to these questions can vary based on the type of

product being considered.

C.   Labeling Considerations

Acceptability of home-use IVD labeling will depend upon

whether such labeling meets the criteria for "adequate

directions for use" as provided for in section

502(f)(1) of the Act and defined by 21 CFR 801.5, and

the requirements of Subpart C of 21 CFR Part 801.  In

addition, home-use IVDs will be deemed to be in

compliance with the requirements for adequate

directions for use if they meet the requirements of 21

CFR 809.10-Labeling for In Vitro Diagnostic Products.

Inherent in the concept of "adequate directions for

use" is the need for the labeling of home-use IVDs to

be simple, concise, easy to understand, make liberal

use of illustrations and drawings, use bold print or

other methods to highlight warnings and precautions,

and provide color coding of reagent containers whenever

practicable.

Labeling for IVDs intended for home and professional

use are covered by 21 CFR 809.10.  The following

elaborates on the application of the regulation to the

home-use situation and suggests means of accomplishing

compliance:

Section 809.10(a)

1.   The intended use statement (section 809.10(a)(2))

should clearly indicate on the outside packaging

container the type of procedure that is offered,

i.e., screening, monitoring, or diagnostic, and

the specific disorder, condition, or risk factor

of interest for which the test is intended.  In

addition, a statement also should be presented

indicating who should use the test and the

conditions for its use, followed by a

contraindication statements), if necessary.

2.   The quantity, proportion, concentration, or

activity of each reactive ingredient included with

the test kit need only be presented as specified

by section 809.10(a)(3) when information is

necessary for proper lay use of the test.

Section 809.10(b)

The order and scope of the information specified by

this section should be as follows:

1.   The intended use statement should reiterate the

information required under section 809 10(a).

2.   Section 809.10(b)(3), summary and explanation of

the test, may be combined with section

809.10(b)(4), principles of the procedure.  This

combination of sections should include a

discussion of both the medical benefits and

limitations of the test with respect to its

intended use in addition to providing a simple

explanation of how the test works.  Presentation

of chemical reactions or formulas are not

necessary but may be referenced in the

bibliography (section 809.10(b)(13)).

3.   A separate information section relating to

interpretation of test results, should be

presented and include information specified by

section 809.10(b)(10), limitations of the

procedure.  This section should indicate the

significance of the test results in light of the

test's intended use and provide specific

information as to what appropriate follow-up

action should be taken by the user. Additionally,

this section should explain the meaning of

false-positive and false-negative test results and

cite possible sources and implications of false

results.

4.   A separate information section should be presented

which identifies any foods, medications, or other

possible interfering substances that would affect

test results. Such information should specify what

substances should be avoided and for how long

prior to testing.

5.   Information specified in sections

809.10(b)(5),(6),(7), and (8) may be combined into

a single information section, i.e., test

procedure.  All information specified by these

individual sections should be presented in the

combined section, if applicable, in addition to

the following:

a.   as previously stated, the quantity,

proportion, concentration, or activity of

each reactive ingredient included in the test

kit need not be presented unless this

information is necessary for proper use of

the test.  However, such information should

be presented elsewhere in a premarket

submission for CDRH review.

b.   discussion of specimen collection,

preparation, and analysis steps should be

enhanced by the use of pictures and

illustrations, preferably in color.

6.   Information specified by section 809.10(b)(9),

test results, should be presented.  This section

also should include trouble-shooting information

where applicable.

7.   Information specified by sections 809.10(b)(11)

and (12) may be combined into a single information

section, i.e., test performance characteristics.

m is section should summarize the data

(separately) from both the laboratory evaluation

and the consumer field evaluation (see III

Premarket Evaluation Considerations) and should

include a concise discussion of the study

protocols used to evaluate the test's performance.

Also, discussion and/or presentation of data

relative to the test's accuracy should be

presented and include false-positive and

false-negative levels.  Such information should be

presented for both the laboratory and the consumer

field evaluations.

8.   Sections 809.10(b)(13), (14), and (15) should be

addressed in the same manner as other IVDs.

In addition, CDRH recommends that manufacturers provide

users with a toll-free telephone number or an address to

write to should questions occur concerning the use of the

device.

Finally, CDRH is concerned that consumers be aware of

significant changes in home-use IVD labeling; particularly

labeling accompanying high frequency use devices such as

home-use glucose test devices. CDRH advises manufacturers to

alert consumers of significant labeling changes by including

special inserts with device packaging which highlight key

labeling changes.  Moreover, CDRH advises manufacturers that

labeling changes that significantly impact on a device's

safe and effective use may require premarket review and

clearance by CDRH before commercial distribution of devices

bearing such labeling.

D.   Premarket Evaluation Considerations

Given the above considerations CDRH recommends that

prospective manufacturers of home-use IVDs conduct the

following performance studies and submit results to

CDRH for review in an appropriate premarket submission:

A.   A laboratory evaluation of the analytical

performance of the device, e.g., analytical

sensitivity, analytical specificity, accuracy, and

reproducibility, should be done by performing

standard laboratory measurements using appropriate

reference tests.  The purpose of this evaluation

is to establish the "true" performance

characteristics of the device as determined under

controlled laboratory conditions; and,

A consumer field evaluation of the device should

be conducted to determine the device's performance

when used by lay users, unassisted, following

instructions provided in the labeling.

CDRH places considerable importance on the outcome of

the consumer field evaluation (item B above) in

assessing the  safety and effectiveness of home-use

IVDs.  Consequently, the following additional

suggestions are offered with respect to the goals,

conduct, and analysis of such studies.

1.   The primary goal of the consumer field evaluation

should be to establish comparable levels of test

performance for a particular home-use IVD when

comparison-tested by both intended users and

trained laboratory personnel.  It is suggested

that split clinical samples be tested by both lay

users and trained laboratory personnel with both

groups using the home-use IVD. Such testing will

simultaneously validate the technical performance

characteristics of the device and the device's

labeling based on user experience.

2.   In order to assess the ability of lay users to

properly perform and interpret test results, CDRH

recommends that manufacturers provide simple

questionnaires to study participants as part of

the consumer field evaluation to determine if lay

users can read and understand the labeling.  For

example, a questionnaire could be used to

determine if the user understands the purpose of

the test, the conditions for its use, the test's

limitations, the meaning of the results and

appropriate follow-up

It is recommended that samples of such

questionnaires be included in a premarket

submission, as well as a tabulation of all study

participants' responses to each question.

3.   Lay users selected for study should be

representative of target users for which a test is

intended.  A representative study population may

require the selection of individuals of varied

background, education levels, and age groups.

Consequently, subjects should be selected for

testing based on carefully developed selection

criteria.  The basis for establishing the

selection criteria should be submitted to CDRH

along with the study protocol and test results.

4.   The number of subjects selected for testing should

be sufficient to substantiate that all targeted

lay user populations can perform and interpret the

test.  It is highly recommended that the number of

subjects selected for study be based on a

statistically valid sampling of relevant lay users

and should take into account appropriate

demographic factors.  Moreover, test results

should be analyzed using appropriate statistical

methods to demonstrate correlation between lay

users and trained technologists performing the

test.

5.   Ideally, to avoid potential problems of bias

during data collection and evaluation of the

consumer field study, a double blind design should

be used.  However, in studies where such a design

is impossible or impractical, a single blind

approach and other measures to reduce potential

bias may be acceptable.  The manufacturer should

be careful to select the appropriate type of study

to ensure that the study will provide appropriate

information.

CDRH believes that both prospective manufacturers of

home-use IVDs and CDRH reviewers will be better able to

evaluate the safety and effectiveness of these devices when

supporting information based on the above considerations is

submitted to CDRH in an appropriate premarket submission.

Moreover, CDRH believes that as a result of these

considerations the general public health will be better

served by the availability of meaningful and reliable

home-use IVDs.

10/5/88