Clinical Evaluation Report

<Product Name>

Evaluators:

<Name, Title>

<Name, Title>

<Name, Title>

Reviewed by:

<Name, Title>

<Name, Title>

Approved by:

<Name, Title>

<Name, Title>

Table of contents

Executive summary	4

1.	Scope of the clinical evaluation	5

1.1	Device description	5

1.1.1	Components	6

1.1.2	Materials	11

1.1.3	Mechanical characteristics	11

1.1.4	Principles of operation	12

1.1.5	Chemical characteristics	12

1.2	Intended purpose of the device	12

1.3	Changes since the last report	14

1.4	Relevant regulations	14

2.	State of the art	16

3.	Identification of relevant clinical data	23

3.1	Literature search strategy	23

3.1.1	Clinical literature search	24

3.1.2	Data selection of clinical literature search results	26

3.1.3	Adverse event database search	27

4.	Summary and appraisal of the clinical data	28

4.1	Criteria for suitability	28

4.2	Criteria for data contribution	29

4.3	Selection criteria	29

5.	Data analysis	31

5.1	Performance	31

5.2	Safety	37

5.3	PMS data	41

5.4	Requirement on acceptable benefit/risk profile	46

5.5	Requirement on acceptability of side effects	47

5.6	Product documentation and instructions for use	47

6.	Conclusions	47

7.	Next clinical evaluation	48

8.	Dates and signatures	49

9.	Qualification of the responsible evaluator	50

10.	Declaration of interests	50

11.	References	51

12.	Appendices	54

LIST OF TABLES

Table 1. List of implants, Inion S-1TM System	7

Table 2. List of implants, Inion S-2TM System	8

Table 3. List of accessory instruments, Inion S-1 and Inion S-2 systems	9

Table 4. Change Log - Clinical Evaluation Inion Spinal Graft Containment System™	14

Table 5. International standards	14

Table 6. European guidance documents	15

Table 7. Comparison of Inion Spinal Graft Containment System™ with equivalent products	22

Table 8. Search queries and search parameters used for literature search in different databases and additional clinical data sources.	24

Table 9. Clinical literature search data selection results	26

Table 10. Suitability criteria	28

Table 11. Criteria for data contribution	29

Table 12. Selection criteria	29

Table 13. Articles included for evaluation	30

Table 14. A systematic review	31

Table 15. Performance related articles	32

Table 16. Safety related articles	37

Table 17. Inion S-1: Sales and risk occurrence rate, long term	41

Table 18. Inion S-2: Sales and risk occurrence rate, long term	42

Table 19. Adverse events from the FDA Maude and TPLC databases (2005-2016).	42

Table 20. Risk occurrence per hazard type, reported adverse events in FDA Maude	44

Table 21. Inion S-1: Hazards analysis of customer feedback evaluations	45

Table 22. Inion S-2: Hazards analysis of customer feedback evaluations	45

Reference

MEDDEV 2.7.1 rev 4 on clinical evaluation

GHTF SG5 N2R8:2007, Appendix D: A Possible Method of Appraisal

- Executive summary

This report presents the clinical evaluation of <Product Name, brief description of main components and accessories>.  <Product Name> consists of <components> made of < description of materials>. The <components 1> include radiographic markers for postoperative imaging, which are made of chemical element Tantalum (Ta) in accordance with ASTM B365 and ASTM F560. The plates are manufactured by compression moulding. The screws are manufactured by extrusion and machining.  This device, in conjunction with traditional rigid fixation, is intended for use in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts.  Inion S-1™ plates and screws are intended for the cervical spine (levels of fixation C2-T1)) and Inion S-2™ plates and screws are intended for the thoraco-lumbar spine (levels of fixation T1 – S1). This device is not intended for load bearing indications.

Inion Spinal Graft Containment System implants retain most of their initial strength up to 16 weeks and gradually lose their strength thereafter. Bioresorption takes place within two to four years. The implants are sterile and non-collagenous. The shelf life of the product is three years. Inion Spinal Graft Containment System implants are designed to be inserted with manual surgical accessory instruments such as drill bits, bone taps, drill guides, plate templates, temporary fixation screw etc. These instruments are provided nonsterile and are reusable.

The clinical evaluation is conducted by collecting and analyzing clinical literature of Inion Spinal Graft Containment System™ and evaluating the PMS data. Clinical data is searched from PubMed, ScienceDirect, Embase, Cochrane Central Register of Controlled Trials (CENTRAL), Clinicaltrails.gov, WHO ICTRP and FDA Manufacturer and User Facility Device Experience (MAUDE) database.

The clinical data analysis concludes that the overall complication rates and risks related to the devices remain continuously low and acceptable.  No clinically relevant change is detected over time, and no new health or safety risks, no new side effects have been discovered during this evaluation.  Anticipated residual risks may occur, but the number is low.

As a result of this clinical evaluation, the evidence provided demonstrates the safety and performance of Inion Spinal Graft Containment System™ implants and instruments in their product-specific indications as described in Instructions for Use, also conformity with the EU Essential Requirements.  It can be concluded that Inion Spinal Graft Containment System™ represents the current state of the art in spinal fusion procedures.

Scope of the clinical evaluation

The objective of this clinical evaluation is to identify, select, review and assess all available clinically relevant data of Inion Spinal Graft Containment System™ implants and accessory instruments.

Conformity assessment with the Medical Devices Directive, MDD 93/42/EEC requires a medical device manufacturer to demonstrate that the claims made in relation to the device’s safety and performance, under the normal conditions of its use, are attainable.  Generally, this requires clinical data (Annex X, MDD), but evidence of the satisfactory clinical safety and performance of a device may be provided in the form of a critical evaluation of published and/or unpublished data on clinical experience with the device, or on a similar device to which equivalence can be demonstrated. This clinical evaluation is submitted to the MDD as amended by directive 2007/47/EEC.

Based on the Essential Requirements and the residual risk findings from the Inion Spinal Graft Containment System™ risk analysis, the scope of this clinical evaluation comes from the intended performance and clinical residual risks in the risk analysis of these products:

1. Bioabsorbable plates and screws as graft containment method for spinal fusion procedures: sufficient strength to maintain relative position of autograft or allograft during spinal fusion and degradation thereafter.

2. Biocompatibility: postoperative tissue reactions in short and long term

3. Issues related to insertion such as implant breakage, instrumentation issues.

Device description

This Clinical Evaluation Report (CER) focuses on Inion Spinal Graft Containment System™ implants and instruments, which have been developed by the project name Inion Spinal Graft Containment System, D358 by Inion Oy, Lääkärinkatu 2, FI-33520 Tampere Finland, and thereafter extended with new products and modifications in subsequent development project Inion S-2 Plate line extension project, D338. The line extension products for Inion S-2 (D338) have not been submitted for CE mark assessment at this time, but have been developed and are in clinical use in other markets.

The first generation Inion S-1 was launched to the market in 2005. The current generation Inion Spinal Graft Containment System™ was CE marked and launched to the market in 2007. Since then, the Inion S-1 and/or Inion S-2 products have been distributed in Europe, USA, and China. Since placing into market, it is estimated that approximately 1500 surgical operations have been conducted with Inion S-1 implants and about 200 operations with Inion S-2 implants. In addition, a similar product family Inion S-1™ Anterior Cervical Fusion system has been marketed since 2005 with an estimated number of 1600 operations up till today.

Components

The Inion Spinal Graft Containment System™ implants include plates and screws offered in different shapes and sizes, which are presented in Tables 1 and 2 below.  The plates and screws are provided sterile by ionizing irradiation.  They are intended for single use and shall not be re-sterilized or re-used. The implants are classified as Class III (93/42/EEC Annex IX Rule 8: Implantable and long-term surgically invasive devices, biodegradable).

The implants are presented in Pictures 1-2.

The Inion Spinal Graft Containment System™ implants are designed to be implanted using specific accessory instrumentation such as drill bits, threading taps, awl, plate holder, drill guides, screw driver, temporary fixation screw, screw removal instrument, plate templates, and a heating device.  The accessory instruments are provided non-sterile and are re-usable.  They are intended to be cleaned, disinfected and sterilized by the user with the steam sterilization. Instructions for cleaning and disinfection of reusable instruments are provided to the user in each instrument package.

The accessory instruments include devices in Class I (93/42/EEC Annex IX Rule 6: Non-sterile, reusable surgically invasive instruments intended for transient use) and Class IIa (Rule 6 and connected to active device).

Picture 1. Inion S-1TM Plate/Screw construct

Picture 2. Inion S-2TM Plate/Screw construct (Original Plate on the left and line extension plate (for US only) on the right)

Table 1. List of implants, Inion S-1TM System

| REF-number | Description |
| --- | --- |
| SPN-5415 | Inion SGCS 120 x 80 x 0.75 mm Plate |
| SPN-5010 | Inion SGCS 21 x 19 x 2.0 mm Plate, 4 holes |
| SPN-5011 | Inion SGCS 23 x 19 x 2.0 mm Plate, 4 holes |
| SPN-5012 | Inion SGCS 25 x 19 x 2.0 mm Plate, 4 holes |
| SPN-5013 | Inion SGCS 27 x 19 x 2.0 mm Plate, 4 holes |
| SPN-5014 | Inion SGCS 29 x 19 x 2.0 mm Plate, 4 holes |
| SPN-5015 | Inion SGCS 31 x 19 x 2.0 mm Plate, 4 holes |
| SPN-5016 | Inion SGCS 33 x 19 x 2.0 mm Plate, 6 holes |
| SPN-5017 | Inion SGCS 36 x 19 x 2.0 mm Plate, 6 holes |
| SPN-5018 | Inion SGCS 39 x 19 x 2.0 mm Plate, 6 holes |
| SPN-5019 | Inion SGCS 42 x 19 x 2.0 mm Plate, 6 holes |
| SPN-5020 | Inion SGCS 45 x 19 x 2.0 mm Plate, 6 holes |
| SPN-5021 | Inion SGCS 48 x 19 x 2.0 mm Plate, 6 holes |
| SPN-5509 | Inion SGCS 2.5 x 8 mm Screw |
| SPN-5220 | Inion SGCS 4.5 x 12 mm Screw |
| SPN-5221 | Inion SGCS 4.5 x 14 mm Screw |
| SPN-5222 | Inion SGCS 4.5 x 16 mm Screw |
| SPN-5223 | Inion SGCS 5.0 x 12 mm Screw |
| SPN-5224 | Inion SGCS 5.0 x 14 mm Screw |
| SPN-5225 | Inion SGCS 5.0 x 16 mm Screw |

Table 2. List of implants, Inion S-2TM System

| REF-number | Description |
| --- | --- |
| SPN-5421* | Plate, 31 mm (16 mm hole span) |
| SPN-5422* | Plate, 34 mm (19 mm hole span) |
| SPN-5423* | Plate, 37 mm (22 mm hole span) |
| SPN-5424* | Plate, 40 mm (25 mm hole span) |
| SPN-5416 | Inion SGCS 120 x 80 x 0.75 mm Plate |
| SPN-5401 | Inion SGCS 37 x 28 x 3.2 mm Plate, 4 holes |
| SPN-5402 | Inion SGCS 40 x 28 x 3.2 mm Plate, 4 holes |
| SPN-5403 | Inion SGCS 43 x 28 x 3.2 mm Plate, 4 holes |
| SPN-5404 | Inion SGCS 46 x 28 x 3.2 mm Plate, 4 holes |
| SPN-5405 | Inion SGCS 49 x 28 x 3.2 mm Plate, 4 holes |
| SPN-5406 | Inion SGCS 50 x 28 x 3.2 mm Plate, 6 holes |
| SPN-5407 | Inion SGCS 53 x 28 x 3.2 mm Plate, 6 holes |
| SPN-5408 | Inion SGCS 56 x 28 x 3.2 mm Plate, 6 holes |
| SPN-5409 | Inion SGCS 59 x 28 x 3.2 mm Plate, 6 holes |
| SPN-5410 | Inion SGCS 62 x 28 x 3.2 mm Plate, 6 holes |
| SPN-5417 | Inion SGCS 70 x 35 x 4.5 mm Plate, 6 holes |
| SPN-5508 | Inion SGCS 2.5 x 8 mm Screw |
| SPN-5500 | Inion SGCS 6.0 x 22 mm Screw |
| SPN-5501 | Inion SGCS 6.0 x 25 mm Screw |
| SPN-5502 | Inion SGCS 6.0 x 28 mm Screw |
| SPN-5503 | Inion SGCS 6.0 x 31 mm Screw |
| SPN-5504 | Inion SGCS 6.7 x 22 mm Screw |
| SPN-5505 | Inion SGCS 6.7 x 25 mm Screw |
| SPN-5506 | Inion SGCS 6.7 x 28 mm Screw |
| SPN-5507 | Inion SGCS 6.7 x 31 mm Screw |

*Note: SPN-5421 SPN-5422 SPN-5423 and SPN-5424 do not carry a CE mark.

Table 3. List of accessory instruments, Inion S-1 and Inion S-2 systems

| Inion S-1 Anterior Cervical Fusion System and Inion Spinal Graft Containment System |  |
| --- | --- |
| REF-number | Description |
| INS-9540 | Drill bit 3.7 mm with stop for 12 mm screw |
| INS-9541 | Drill bit 3.7 mm with stop for 14 mm screw |
| INS-9542 | Drill bit 3.7 mm with stop for 16 mm screw |
| INS-9543 | Drill bit 4.2 mm with stop for 12 mm screw |
| INS-9544 | Drill bit 4.2 mm with stop for 14 mm screw |
| INS-9545 | Drill bit 4.2 mm with stop for 16 mm screw |
| INS-9546 | Bone tap 4.5 mm with stop for 12 mm screw |
| INS-9547 | Bone tap 4.5 mm with stop for 14 mm screw |
| INS-9548 | Bone tap 4.5 mm with stop for 16 mm screw |
| INS-9549 | Bone tap 5.0 mm with stop for 12 mm screw |
| INS-9550 | Bone tap 5.0 mm with stop for 14 mm screw |
| INS-9551 | Bone tap 5.0 mm with stop for 16 mm screw |
| INS-9552 | Drill bit 2.0 mm |
| INS-9553 | Screw removal instrument |
| INS-9554 | Drill guide, single |
| INS-9555 | Drill guide, twin |
| INS-9556 | Plate holder |
| INS-9527 | Temporary fixation screw |
| INS-9558 | Screw driver for temporary fixation screw |
| INS-9559 | Screw driver shaft |
| INS-9560 | Awl |
| ACC-9841 | Sterilization tray for instruments |
| ACC-9823 | Plate template 21 x 19 mm |
| ACC-9824 | Plate template 23 x 19 mm |
| ACC-9825 | Plate template 25 x 19 mm |
| ACC-9826 | Plate template 27 x 19 mm |
| ACC-9827 | Plate template 29 x 19 mm |
| ACC-9828 | Plate template 31 x 19 mm |
| ACC-9829 | Plate template 33 x 19 mm |
| ACC-9830 | Plate template 36 x 19 mm |
| ACC-9831 | Plate template 39 x 19 mm |
| ACC-9832 | Plate template 42 x 19 mm |
| ACC-9833 | Plate template 45 x 19 mm |
| ACC-9834 | Plate template 48 x 19 mm |
| Inion S-2 Inion S-2 Anterior Thoraco-Lumbar Fusion System |  |
| REF-number | Description |
| INS-9561 | Drill bit 5.0 mm with stop for 22 mm screw |
| INS-9562 | Drill bit 5.0 mm with stop for 25 mm screw |
| INS-9563 | Drill bit 5.0 mm with stop for 28 mm screw |
| INS-9564 | Drill bit 5.0 mm with stop for 31 mm screw |
| INS-9565 | Drill bit 5.7 mm with stop for 22 mm screw |
| INS-9566 | Drill bit 5.7 mm with stop for 25 mm screw |
| INS-9567 | Drill bit 5.7 mm with stop for 28 mm screw |
| INS-9568 | Drill bit 5.7 mm with stop for 31 mm screw |
| INS-9569 | Bone tap 6.0 mm with stop for 22 mm screw |
| INS-9570 | Bone tap 6.0 mm with stop for 25 mm screw |
| INS-9571 | Bone tap 6.0 mm with stop for 27 mm screw |
| INS-9572 | Bone tap 6.0 mm with stop for 31 mm screw |
| INS-9573 | Bone tap 6.7 mm with stop for 22 mm screw |
| INS-9574 | Bone tap 6.7 mm with stop for 25 mm screw |
| INS-9575 | Bone tap 6.7 mm with stop for 28 mm screw |
| INS-9576 | Bone tap 6.7 mm with stop for 31 mm screw |
| INS-9577 | Plate holder, straight |
| INS-9578 | Plate holder, angled |
| INS-9579 | Drill guide, 6.0 mm |
| INS-9580 | Awl |
| INS-9581 | Temporary fixation screw |
| INS-9582 | Screwdriver shaft for temporary fixation screw |
| INS-9583 | Screwdriver shaft |
| INS-9585 | Drill guide, 6.7 mm |
| ACC-9842 | Sterilization tray for instruments |
| ACC-9843 | Plate contouring block |
| ACC-9845 | Plate template 37 x 28 mm |
| ACC-9846 | Plate template 40 x 28 mm |
| ACC-9847 | Plate template 43 x 28 mm |
| ACC-9848 | Plate template 46 x 28 mm |
| ACC-9849 | Plate template 49 x 28 mm |
| ACC-9850 | Plate template 50 x 28 mm |
| ACC-9851 | Plate template 53 x 28 mm |
| ACC-9852 | Plate template 56 x 28 mm |
| ACC-9853 | Plate template 59 x 28 mm |
| ACC-9854 | Plate template 62 x 28 mm |
| C60 | Plate holder |
| C61 | Drill Guide |
| C63 | Screwdriver attachment block |
| ACC-9855 | Template, 31 mm (16 mm hole span) |
| ACC-9856 | Template, 34 mm (19 mm hole span) |
| ACC-9857 | Template, 37 mm (22 mm hole span) |
| ACC-9858 | Template, 40 mm (25 mm hole span) |
| ACC-9859 | Template, 31-34 mm (16-19 mm hole span) |
| ACC-9860 | Template, 37-41 mm (22-25 mm hole span) |
| Inion Spine Universal instruments |  |
| REF | Description |
| INS-9024 | Plate bending pliers |
| INS-9120 | Screwdriver handle, cannulated |
| INS-9121 | Screwdriver handle, ratchet |
| ACC-9802 | Inion Thermo drape |
| ACC-9840 | Inion Thermo+ (water bath 110 V + US plug) |
| ACC-9810 | Inion Thermo+ (water bath 230 V + EU plug) |

Materials

Inion Spinal Graft Containment System™ consists of biodegradable plates and screws made of resorbable polylactic acid copolymer P(L/DL)LA 80:20. These polymers have a long history of safe medical use and they degrade in vivo by hydrolysis into alpha-hydroxy acids that are metabolized by the body (Ambrose et al. 2004; Gunatillake et al. 2003; Middleton et al. 2000). The implants include radiographic markers for postoperative imaging, which are made of chemical element Tantalum (Ta) in accordance with ASTM B365 and ASTM F560. Tantalum markers have a long history of safe medical use (Levine et al. 2006b).

The plates are manufactured by extrusion and compression moulding. The screws are manufactured by extrusion and machining.

The surgically invasive parts of Inion Spinal Graft Containment System™ accessory instruments are manufactured from surgical grade stainless steel in accordance with ASTM F899, or titanium.

Mechanical characteristics

The Inion Spinal Graft Containment System™, in conjunction with traditional rigid fixation, is intended for use in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts and autografts. The subsystem Inion S-1™ Biodegradable Anterior Cervical Fusion System plates and screws are intended to be used in the cervical spine, levels C2-T1, and the subsystem Inion S-2™ Biodegradable Anterior Thoraco-Lumbar Fusion System is intended to be used in the thoraco-lumbar spine, levels T1-S1. This device is not intended for load bearing indications.

Mechanical properties of plates and screws have been tested and verified during product development initially in push-off, tensile, torsion, lateral bending and cyclic compressive bending tests.  During in vitro degradation, the mechanical properties of the plates have been determined and verified by tensile test and the properties of the screws by a shear test. All test reports and test data are filed at Inion Oy, quality management system.

Principles of operation

The implants will be used in internal fixation to maintain the relative position of weak bony tissue such as allografts or autografts in spinal fusion procedures, in order to prevent graft (and/or fusion cage) extrusion/dislodgement induced complications after spinal fusions.  Based on in vitro testing, the implants retain most of their initial strength up to 16 weeks and gradually lose their strength thereafter; and bioresorption takes place within two to four years. The plates SPN-5010 – SPN-5021, SPN-5401 – SPN-5410, SPN-5417 and screws SPN-5220 – SPN-5225, SPN-5500 – SPN-5507 contain radiopaque tantalum markers for postoperative radiographic imaging.  Plates are prepared by soaking them in heated sterile water, to become temporarily malleable and thereby can be shaped as desired.

The main steps of the surgical operation with Inion Spinal Graft Containment System™ are:

-  Selection of appropriate implant size

-  Achieve malleability of the plates by soaking in Inion Thermo+™ water bath

-  Drilling and tapping the bone hole

-  Fastening of the plates with screws.

Chemical characteristics

Biological safety of the Inion Spinal Graft Containment System™ has been evaluated and verified by Inion in accordance with ISO 10993-serie standards including tests for cytotoxicity and chemical analysis (quantification of heavy metals). Test reports are filed at Inion Oy, quality management system.

Intended purpose of the device

Indications:

Inion Spinal Graft Containment System, in conjunction with traditional rigid fixation, is intended for use in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts. This device is not intended for load bearing indications.

Contraindications:

Inion Spinal Graft Containment System implants are not intended for use in, and are contraindicated for:

• Active or potential infection

• Cancer

• Pseudoarthrosis

• Patient conditions including limited blood supply, insufficient quantity or quality of bone; and where patient cooperation cannot be guaranteed (e.g., alcoholism, drug abuse).

Precautions:

- Instruments are available to aid accurate implantation of the Inion Spinal Graft Containment implants. Surgical instruments are subject to wear with normal usage and may break. Surgical instruments are only to be used for their intended purpose. All instruments are to be regularly inspected for wear and damage. Use only the Inion Spinal Graft Containment System instruments.

- Use only Inion Spinal Graft Containment screws to secure the Inion Spinal Graft Containment plates. DO NOT use Inion Spinal Graft Containment screws to secure plates or sheets other than those of Inion® (e.g., metal plates or sheets).

- DO NOT use for unintended applications! The proper function (i.e., efficacy and safety) of these implants cannot be guaranteed in the event of off-label use.

- The safety and effectiveness of this device, as an adjunct to fusion, when used without rigid supplemental internal fixation has not been established. This device is not designed to withstand physiologic loads when used by itself.

- The patient should be warned that the implants can break or loosen as a result of early stress, activity or load bearing. Premature discontinuation of additional external immobilization (i.e., collar) may cause non-union or mal-union.

- The surgeon should consider the levels of implantation, patient weight, patient activity level, other patient conditions, etc. which may impact on the performance of the system.

- The implantation of pedicle screw spinal systems should be performed only by experienced spinal surgeons with specific training in the use of the specific pedicle screw spinal system because the procedure is technically demanding and presents a risk of serious injury to the patient.

Warnings:

- The Inion Spinal Graft Containment System devices provide fixation and are not intended to replace healthy tissues or withstand the stress of full load bearing.

- Incorrect selection, placement, positioning, or fixation of the implant can cause subsequent undesirable results or breakage of implants or instruments. The surgeon should be familiar with the devices, the method of application and the surgical procedure prior to performing the surgery.

- The safety and efficacy of pedicle screw spinal systems have been established only for spinal conditions with significant mechanical instability or deformity requiring fusion with instrumentation. These conditions are significant mechanical instability or deformity of the thoracic, lumbar, and sacral spine secondary to severe spondylolisthesis with objective evidence of neurological impairment, fracture, dislocation, scoliosis, kyphosis, spinal tumor, and failed previous fusion (pseudoarthrosis). The safety and efficacy of pedicle screw spinal systems for any other conditions are unknown.

Changes since the last report

Table 4. Change Log - Clinical Evaluation Inion Spinal Graft Containment System™

| Version | Date / Year | Changes |
| --- | --- | --- |
| 1.0 | JUN 2007 | Initial report |
| 2.0 | APR 2012 | Update acc. to MEDDEV 2.7/1 rev. 3 |
| 3.0 | JUN 2013 | Update |
| 4.0 | MAR 2014 | Update |
| 5.0 | APR 2015 | Update |
| 6.0 | APR 2017 | Update acc. to MEDDEV 2.7/1 rev. 4 |

Changes to the product

No changes to the product composition or manufacturing have been introduced since the last clinical evaluation (Version 5.0, see Table 4). The Inion S-2 product portfolio was extended with 2-hole plates in 2015. These implants are currently marketed only outside the EU and do not carry a CE-mark.

Relevant regulations

European Legislation

Council Directive 93/42/EEC concerning medical devices, last amended by Directive 2007/47/EC (MDD)

Table 5. International standards

| EN ISO 13485 (2012) | Medical devices — Quality management systems-Requirements for regulatory purposes (ISO 13485:2003 + Cor.1:2009) |
| --- | --- |
| EN 556-1 (2001) | Sterilization of medical devices - Requirements for medical devices to be designated "STERILE" - Part 1: Requirements for terminally sterilized medical devices |
| EN ISO 11607-1 (2009) | Packaging for terminally sterilized medical devices - Part 1: Requirements for materials, sterile barrier systems and packaging systems |
| EN ISO 11607-2 (2006) | Packaging for terminally sterilized medical devices - Part 2: Validation requirements for forming, sealing and assembly processes |
| EN ISO 11137 – 1 (2015) | Sterilization of health care products - Radiation - Part 1: Requirements for development, validation and routine control of a sterilization process for medical devices |
| EN ISO 11137-2 (2015) | Sterilization of health care products - Radiation - Part 2: Establishing the sterilization dose |
| EN ISO 11737-1 (2006) | Sterilization of medical devices - Microbiological methods - Part 1: Determination of a population of microorganisms on products |
| EN ISO 11737-2 (2009) | Sterilization of medical devices - Microbiological methods - Part 2: Tests of sterility performed in the definition, validation and maintenance of a sterilization process |
| EN ISO 14971 (2012) | Medical devices-Application of risk management to medical devices (ISO 14971:2007, Corrected version 2007-10-01) |
| EN ISO 14155 (2012) | Clinical investigation of medical devices for human subjects Good clinical practice (ISO 14155:2011 + Cor. 1 :2011) |
| EN ISO 10993 | Biological evaluation of medical devices |
| EN 1041 (2013) | Information supplied by the manufacturer of medical devices (EN 1041:2008+A1:2013) |
| EN 980:2008 | Symbols for use in the labelling of medical devices |

European guidance documents

Table 6. European guidance documents

| MEDDEV 2.4/1 rev.9 (2010) | Medical devices: Guidance document: Classification of medical devices |
| --- | --- |
| MEDDEV 2.7.1 rev.4 (2016) | Clinical evaluation: A guide for manufacturers and notified bodies under directives 93/42/EEC and 90/385/EEC |
| MEDDEV 2.12-2 rev 2 (2012) | Guidelines on post market clinical follow up |

State of the art

Spinal fusion procedure is a widely accepted treatment for degenerative or traumatic instability of the spine. To reduce or eliminate complications such as implant migration and failure, imaging interference, and stress shielding that are occasionally associated with spinal instrumentation, attention has been given to the use of biodegradable devices. Unlike metal, the elastic modulus of the biodegradable materials compare favourably to bone. The potential advantages of biodegradable materials in spine surgery are significant and include no interference with radiographic studies, predictable degradation through bulk hydrolysis as carbon dioxide and water, and the elimination of long-term residual metal hardware. Future designs of biodegradable fixation devices may begin to resemble the metallic counterparts. Concerns related to design will be strength, design, ability to identify presence and location radiographically, and ease of application and revision.

The spine is a strong and flexible column made up of a series of bones (vertebrae), discs and ligaments. The muscles of the spine are important in maintaining its stability and balance as well as assisting in spinal movement. The cervical or neck region is normally made up of 7 bones, the thoracic or chest region has 12 bones, and the lumbar or low back region is made up of 5 bones. The sacral region consists of 5 fused bones, and the coccygeal region 3 to 5 tiny bones. The vertebrae provide the anterior support and structure of the spine. Each spinal segment is subject to a variety of loads and forces. These include loads from the weight of body segments as well as from the weights being handled. The pattern of injury to the spine is a function of the type and size of forces present at time of injury. There are four types of forces acting on spinal structures: compression, tension, shear, and torsional. The discs are in between the vertebrae and act like a “shock absorber” between the vertebrae. The discs also contribute to the flexibility and mobility of the spinal column.

Disc material is mainly composed of water and proteins. As a normal part of aging, the water content gradually reduces. This can cause the disc to flatten out and even develop tears or cracks throughout the outer part of the disc (anulus fibrosus). These discs are called degenerative discs and they may or may not cause pain. In the case of a degenerative disc, the inner jelly material of the disc (nucleus pulposus) can bulge out and press up against the anulus fibrosus. This can cause pain. Also the cracks or tears of the anulus fibrosus themselves can become a source of pain. Finally, the inner nucleus can also come out through the cracks in the anulus and compress nerves or spinal cord, a condition that may cause weakness, pain, pins and needles or numbness, and may require surgery.

Due to the above described degeneration, the vertebrae get closer together and cause nerve irritation, which usually stems from a ruptured disc, bone spurs, or stenosis. Some herniated discs (i.e., ruptured or slipped discs, or herniated nucleus pulposus) occur from injuries or sudden movements: most arise spontaneously and often occur at night while sleeping. With the aging wear and tear of the spine, some patients develop bony outgrowths. These growths are bone spurs, also known as osteophytes. Bone spurs are the body’s natural response to the inflammation that results from the aging spine. The collection of calcium that turns into the bone spur is a type of natural fusion. However, as they grow and extend, the vertebral openings become narrow. Either the spinal canal and/or the foramen, the opening for nerve passageways, become smaller. This narrowing is stenosis, and results in a pinching (compression) of the spinal cord or the spinal nerve root. Symptoms include pain, weakness, numbness and loss of coordination in the neck or extremities. Movement (vertebral motion) causes the chronic pain.

Conservative non-surgical treatment options for people with disc protrusions in the spine include rest, heat, pain medications and physiotherapy. When non-surgical treatment options fail, surgery is often the next step. This usually means spinal fusion surgery. Pain with compression of the nerves in spinal cord is a common condition that affects the spine and may require surgery. If the spinal cord is being compressed, surgery to decompress the spinal cord is usually recommended. This neurosurgical procedure is performed to relieve the pressure on one or more nerve roots, or on the spinal cord. It involves the stabilization of two or more vertebrae by locking them together (fusing them). The fusion stops the vertebral motion and as a result, the pain is stopped. The compression can be caused by bulging disc or bony spurs.

Surgery on the spinal cord is performed either from the back (laminectomy) or through the front of the spinal cord (discectomy or vertebrectomy). If the compression of the spinal cord is from the front, then the decompression must be done from the front (anterior decompression). Typically, if the entire disc is removed, it is replaced with a wedge of bone (autograft from patient’s own hip or allograft from bone bank, a plug of bone that allows the fusion of bone to happen), and/or synthetic bone graft substitute (e.g., bioactive glass, tricalcium phosphate or hydroxyapatite), and/or a fusion cage/spacer, and thereafter anterior and/or posterior instrumentation (i.e., traditional rigid fixation) is applied for stabilization. Sometimes a plate and screws (i.e., anterior instrumentation) are used to hold the graft in place (graft containment). When the procedure is carried out through the front of the spinal cord it is called anterior discectomy and fusion.

Anterior surgery for the removal of herniated cervical discs was first used in the 1950s with the introduction of the Smith-Robinson and Cloward procedures. These techniques have become the golden standard for a variety of indications in the spine such as degenerative disease, tumor, inflammatory process, trauma and failed fusion. Although the advantages of anterior plate stabilization have been reported (i.e. immediate stability, restoration of lordotic curve, shortened fusion time and enhanced fusion quality) some clinicians have doubted whether this is truly beneficial to the patient or rather over treatment. However, it has been shown in many scientific reports that anterior plating in degenerative disc disease does not constitute over treatment but rather supplements the internal stabilization initially provided by the bone graft and yields a lower reoperation rate (Alvarez et al. 1999, Caspar et al. 1998). The goals of interbody bone graft placement and subsequent fusion are to improve and/or to maintain interbody height, to establish and/or to maintain physiological spinal geometry, and to achieve arthrodesis to eliminate pathological motion (Geer et al. 1999). The problem with fusing bones in the spine is that adjacent levels in the spine are placed under more strain and therefore a person that has a single level fusion may have to have another operation for narrowing at the next level either above or below. Secondly, the bone does not always heal or fuse correctly (e.g. in case of graft extrusion/dislodgement).

Spinal fusion procedure is a widely accepted treatment for degenerative or traumatic instability of the spine. Spinal graft containment devices (plates, sheets, etc.) that maintain the relative position of weak bony tissue such as allografts or autografts can be used to prevent graft (and/or fusion cage) extrusion/dislodgement induced complications after spinal fusions. Both non-degradable (i.e., metal plates or meshes) and degradable products (i.e. bioabsorbable sheets/plates) are available.

Despite many improvements, the use of metallic spinal implants is not without complications. Complications associated with the placement of metallic plates have a reported incidence as high as 35% (Lowery et al. 1998), and include screw backout, graft extrusion, and dysphagia and odynophagia when used in the cervical spine. Typically, the failure can occur in the form of screw back-out or breakage with subsequent displacement of the implant, and extrusion of the bone graft. In case of implantation in the cervical spine, dislodgement can damage the esophagus, and if the gastrointestinal tract is violated implant can cause multi-organ failure (Lamesch et al. 1997). There have been several case reports of vessel injuries (Knoop et al. 1999), patients coughing up screws (Geyer et al. 2001) or bone grafts (Cavanagh et al. 1996) into the oral cavity, and one report of a plate disappearing (Fujibayashi et al. 2000). In addition, metallic plates interfere with imaging of the fusion mass via x-rays, CT or MRI scanning, they may serve as a locus for bacterial infection, and may be physiologically (and even psychologically) taxing to many patients. Any metallic plate implanted in orthopaedic surgery may necessitate a second surgery for plate removal, and in case of spinal implantation, the possibly needed removal procedure is more challenging and expensive than with more superficially placed implants (e.g., in the treatment of fractures in the extremities).

Because of these complications, consideration for the application of biodegradable implant in this technology has gained momentum. Biodegradable graft containment sheets and plates in spine fusion may avoid many of the disadvantages related with the use of metallic devices. Biodegradable implants degrade uniformly in a biological environment. Their breakdown products incorporate into normal cellular, physiologic, and biomechanical processes. These materials have been shown to be biocompatible with degradation products that are tolerated, with no immunogenic or mutagenic tendencies. Foreign body reactions are the most frequently reported complications in the use of bioabsorbable implants. (Wuisman et al. 2006) In a study of Böstman et al. (2000) 2528 patients underwent various reconstructive procedures with the use of bioresorbable implants. At follow-up, 108 patients (4.3%) were noted to have a clinically significant local inflammatory foreign body reaction. However, it should be noted that foreign body reactions have also been reported with the use of metallic spinal implants. One study reported 1.7% of 174 patients suspected to have a late infection actually had a foreign body reaction to the spinal implant and they were treated by implant removal (Aydinli et al. 1999). Generally, the incidence of infections has been shown to be comparable in patients with resorbable implants compared with metal implant devices. According to Vaccaro et al. (2003), Andrews et al. (2000) evaluated the incidence of infection in a series of 296 patients in which 146 received craniotomy fixation with titanium implants and 150 received craniotomy fixation with a resorbable poly- D,L-lactic acid copolymer. Forty-three patients in the titanium group and 37 patients in the polymer group received postoperative irradiation. The incidence of infection in the titanium and resorbable polymer groups was 4.6% and 4.0%, respectively. In 2007, Bindal et al described the first series of patients to undergo ACDF with the first commercially available resorbable cervical plate for single-level degenerative disc disease. Healing in kyphosis and pseudarthrosis was significantly more likely with resorbable implants as compared with the use of titanium plating. (Bindal et al. 2007) Further, Whitmore et al. (2011) reported a novel application of a bioresorbable plate in the management of a 69-year-old female with multi-level, long-standing cervical spondylotic myelopathy. The patient underwent a revision anterior fusion using a bioresorbable cervical plate and a fibular allograft, in conjunction with posterior fusion with metal instrumentation. Successful fusion is visualized on CT imaging at 8-year follow-up.

The main benefit of a resorbable material is that they confer initial and intermediate stability without having any of the long-term complications, such as stress shielding. Flexible and less rigid bioabsorbable implants confer the advantage of stabilization the motion segments while, over time, allowing a greater transfer of load to the healing bone during implant resorption, potentially minimizing junctional degeneration (Robbins et al. 2004). Other potential advantages of resorbable implants in spinal surgery include a more physiological modulus of elasticity in relation to native vertebral bone allowing more optimal load-sharing, as compared with rigid metal devices, and the elimination of long-term residual metal hardware with the potential attendant complications of migration, dislodgement or hardware removal (Vaccaro et al. 2003). Biodegradable implants allow optimal postoperative radiographic evaluation because of their radiolucency and the absence of artefacts associated with similar metallic devices exposed to advanced imaging modalities (McAfee et al. 2001). However, the postoperative position of implanted biodegradable implants can be evaluated radiographically if the implants contain radiopaque markers. In addition, contrary to metal devices, bioabsorbable implants can be easily cut and contoured to the desired shape at operation and their transparency enables perfect visualisation of the bone graft and the vertebrae during fixation of the implants. However, the biodegradable sheets/plates are not intended to be used alone in high load bearing indications and must therefore be used in conjunction with traditional rigid fixation in the spine (i.e., the biodegradable sheet/plate maintains the relative position of the bone graft and traditional rigid fixation is used to provide stabilization and load-bearing). Future designs of biodegradable fixation devices may begin to resemble the metallic counterparts. Concerns related to design will be strength, design, ability to identify presence and location radiographically, and ease of application and revision.

Tantalum markers

Tantalum has been used extensively in neurosurgery for more than 50 years (Smith 1973). One of the first scientific reports on the use of tantalum in repair of congenital bony defects of skull and spine was reported in 1951 by Scoville and McLaurin. Tantalum has excellent biocompatibility and is safe to use in vivo as evidenced by its historical and current use in pacemaker electrodes, cranioplasty plates and as radiopaque markers. The bioactivity and biocompatibility of tantalum derives from its ability to form a self-passivating surface oxide layer. This surface layer leads to the formation of a bone-like apatite coating in vivo and affords excellent bone and fibrous in-growth properties allowing for rapid and substantial bone and soft tissue attachment. (Levine et al. 2006b)

Tantalum is known to be relatively inert in vivo and is now used as plates, sutures, radio-markers, and prostheses covering a wide variety of medical subspecialties and procedures. For example, Biomet Tantalum Beads – Radiographic Marker (K010348) are used as radiopaque markers and may be implanted into bone or soft tissue. These devices are used to measure movement of implants after surgery with the aid of an X-ray. Implant surgery associated with the use of radiographic markers may include total joint replacement, soft tissue repair, and bone fixation.  Bobyn et al. (1999a) reported good biocompatibility of tantalum in a canine joint replacement model in a long-term follow-up. Dental implants have yielded excellent 8-year follow-up with tantalum-based components used for osseous anchorage according to Bobyn et al. (1999b). The oxide formed on the surface of tantalum implants (self-passivation) in vivo has been found to be quite stable over a wide range of pH and potential ranges (Kato et al. 2000). Matsuno et al. (2001) implanted tantalum wire into soft tissues and the femora of rats in order to determine soft and hard tissue reactions to this metal. The authors found tantalum to have good biocompatibility without a surrounding inflammatory response.

There are several reports where permanent metallic neurosurgical implants have been tested for tissue reactions. The common opinion is that tantalum shows good biocompatibility. In the study of McFadden (1972) the author compared 18 most common metal devices used as permanent neurosurgical implant in dogs for 6 months. The results showed clearly that Tantalum caused the least tissue reactions in the implanted sites (brain, spine, carotids, rectus).

References indicate that there is no known human disease which is attributable to tantalum, that systemic poisonings in industrial situations are unknown, and that tantalum and tantalum compounds are not listed as presumptive or possible carcinogens. The oral LD50 for tantalum pentoxide in rats is quoted to be greater than 8 g/kg bodyweight. Where labelled tantalum has been injected into animal models, only 15% is retained within the body, the balance being rapidly excreted. 40% of that which is retained within the body is retained within bone.

The Inion Spinal Graft Containment System™ consists of plates, meshes and screws made of degradable co-polymers composed of L-lactic and D-lactic acid. These polymers have a long history of safe medical use and they degrade in vivo by hydrolysis into alpha-hydroxy acids that are metabolised by the body. The plates and screws contain radiopaque tantalum markers for postoperative imaging. The MacroPore Hydrosorb Spine System implants include radiopaque barium sulfate beads, which are also non-degradable. However, the biocompatibility of barium sulfate as a radiopaque marker in orthopedic devices can be considered to be somewhat controversial (Cenni et al. 2002, Chang 1981, Sabokbar et al. 1997, Togawa et al. 2006, Wimhurst et al. 2001) compared to the biocompatibility of tantalum. Bioabsorbable plates are typically made malleable before placing them on the patient by soaking the plates in heated water. These techniques utilize the thermo forming properties of the polymers.

Inion Spinal Graft Containment System™ implants retain most of their initial strength up to 16 weeks and gradually lose their strength thereafter. During biomechanical and mechanical comparison testing with Macropore OS Spinal System/Macropore Hydrosorb implants and the Inion Spinal Graft Containment System™ implants, the results were found to be statistically equivalent.

Table 7 shows a comparison of main characteristics of Inion Spinal Graft Containment System™ and equivalent products. Based on the evaluation it can be concluded that Inion Spinal Graft Containment System™ represents the current state of the art in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts.

Table 7. Comparison of Inion Spinal Graft Containment System™ with equivalent products

| Property | Inion Spinal Graft Containment System | Macropore OS Spinal System (K010911) | MacroPore Hydrosorb Spine System (K041105) | Inion S-1TM Biodegradable Anterior Cervical Fusion System |
| --- | --- | --- | --- | --- |
| Intended use | In conjunction with traditional rigid fixation, intended for use in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts. The device is not intended for load bearing indications. | In conjunction with traditional rigid fixation, intended for use in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts. The device is not intended for load bearing indications. | In conjunction with traditional rigid fixation, intended for use in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts. The device is not intended for load bearing indications. | In conjunction with adjunctive immobilisation (i.e. simultaneous conservative treatment with a collar), is intended for use in anterior cervical discectomy and fusion procedures as a means to maintain the relative position of block bone allografts or autografts. This device is not intended for use in other areas of the spine, and is not for load bearing indications. |
| Screw diameter/ Screw length | 2.5 mm – 6.7 mm 8 mm – 31 mm | 2.0 mm – 4.8 mm | 2.4 mm – 6.5 mm 4 mm – 30 mm | 4.5 mm – 5.0 mm 12 mm – 16 mm |
| Plate thickness Plate size | 0.75 – 4.5 mm 21 x 19 – 120 x 80 mm | 0.75 mm – 3.0 mm 50 x 50 – 120 x 120 mm | 1.0 – 5.0 mm 24 x 18 – 100 x 100 mm | 0.8 mm – 2.0 mm 25 x 19 – 31x 19 mm |
| Marker | Nondegradable Tantalum sphere, diameter: 1.0 mm | -- | Nondegradable Barium sulphate bead, diameter: 0.5 – 2.0 mm | --- |
| Polymeric Material | P(L/DL)LA | P(L/DL)LA | P(L/DL)LA | P(L/DL)LA |
| Degradation | The implants retain most of initial strength 16 weeks and gradually lose strength thereafter. Bioresorption takes place within two to four years. | Retains approximately 80% of initial strength for 12 weeks, 60% at 6 months | Retains approximately 80% of initial strength for 12 weeks, 60% at 6 months | The implants retain most of initial strength 16 weeks and gradually lose strength thereafter. Bioresorption takes place within two to four years. |
| Sterilization | gamma irradiation | gamma irradiation | gamma irradiation | gamma irradiation |

Identification of relevant clinical data

The purpose of this clinical evaluation is to identify and evaluate sufficient clinical data from published clinical studies and scientific literature, to demonstrate the safety and performance of the Inion Spinal Graft Containment in their intended use. In case sufficient clinical data is available on the devices themselves, data search is not needed to be extended to equivalent devices (see Table 7).

In addition, the post-market data of Inion Spinal Graft Containment System™ is identified, appraised, and analyzed.

The clinical data search, appraisal and analysis were performed according to the following procedure presented in Picture 3.

Picture 3. Literature search, appraisal and analysis procedure

Literature search strategy

The following scientific databases were chosen as literature search sources:

- Pubmed

- ScienceDirect

- Embase

- Cochrane CENTRAL

In addition, clinical data was searched from

- Clinicaltrials.gov

- WHO ICTRP

- Citations referenced in the searched literature

- Articles identified in previous Inion Spinal Graft Containment System™ clinical evaluation literature searches

Furthermore, adverse event data was searched from FDA Manufacturer and User Facility Device Experience (MAUDE) and Total Product Life Cycle (TPLC) databases.

Clinical literature search

The search queries and search parameters used for literature search in different databases and additional clinical data sources are summarized in Table 8:

Table 8. Search queries and search parameters used for literature search in different databases and additional clinical data sources.

| Database | Search query | Search timeframe | Search date | Number of search results | Number of search results excluded before data selection | Number of search results included into data selection | Raw data appendix number |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PubMed | inion AND (s-1 OR s1 OR s-2 OR s2); inion AND (plate OR plating OR plates) AND (spine OR spinal); inion AND (cervical OR thoracic OR lumbar OR thoraco-lumbar); (biodegradable OR bioabsorbable OR bioresorbable OR degradable OR absorbable OR resorbable) AND (spine OR spinal) AND (plate OR plating OR plates); (biodegradable OR bioabsorbable OR bioresorbable OR degradable OR absorbable OR resorbable) AND (cervical OR thoracic OR lumbar OR thoraco-lumbar OR fusion) AND (plate OR plating OR plates) | 2005/01/01 –2017/12/31 | 2017/02/20 | 78 | 24 | 54 | Appendix 1 |
| Science Direct | See above | 2005/01/01 –2017/12/31 | 2017/02/20 | 19 | 7 | 12 | Appendix 2 |
| Embase | inion:df or inion:dn | Not limited | 2017/02/14 | 60 | 5 | 55 | Appendix 3 |
| Cochrane Central Register of Controlled Trials (CENTRAL) | inion (All text) | Not limited | 2017/02/14 | 31 | 3 | 28 | Appendix 4 Appendix 5 |
| Clinical trials.gov | inion | Not limited | 2017/02/27 | 16 | 0 | 16 | Appendix 6 |
| WHO ICTRP | spine OR spinal / “condition” AND spinal graft containment or S-1 or s 1 or s-2 or s2 / “intervention” | Not limited | 2017/02/27 | 9 | 0 | 9 | Appendix 6 |
| Citations referenced in the searched literature | - | - | - | 0 | 0 | 0 | - |
| Articles identified in previous clinical evaluations | - | - | - | 6 | 5 | 1 | Appendix 7 |

The search queries were defined with the goal of finding a reasonable number of pertinent data, which would either directly demonstrate adequate clinical performance and clinical safety of the device (i.e pivotal data), or provide an indirect supportive role.

The timeframe used in PubMed and ScienceDirect literature searches was 2005 - present, because the Inion S-1 has been on the market since the year 2005. The timeframe used in Embase and Cochrane searches was not limited.

Results of literature searches of PubMed, Science Direct, Embase and Cochrane were exported from the databases to an Excel, and at that point a sanity check was made to see if there were any duplicates or data entries that were incomplete and which needed to be excluded before data selection.

Data selection of clinical literature search results

The search results from PubMed, Science Direct, Embase and Cochrane were exported into an Excel. Also, Citations referenced in the searched literature and Articles identified in previous Inion Spinal Graft Containment System™ clinical evaluations were evaluated in the Excel. A data selection was made by reviewing the publications’ title and/or abstracts. Publications were excluded if the studied products clearly did not include Inion Spinal Graft Containment System™; a language other than English was used; the article in question did not examine humans; or no clinical data was available. These exclusions were marked as “EBA” (i.e. Excluded Before Appraisal) in the data selection and appraisal excel file (see Appendix 7).

Please see the following Table 9 for a summary of data selection results:

Table 9. Clinical literature search data selection results

| Database | Search query | Number of search results included into data selection | Number of search results excluded (EBA) after data selection | Number of search results included (INC) after data selection for data appraisal |
| --- | --- | --- | --- | --- |
| PubMed | inion AND (s-1 OR s1 OR s-2 OR s2); inion AND (plate OR plating OR plates) AND (spine OR spinal); inion AND (cervical OR thoracic OR lumbar OR thoraco-lumbar); (biodegradable OR bioabsorbable OR bioresorbable OR degradable OR absorbable OR resorbable) AND (spine OR spinal) AND (plate OR plating OR plates); (biodegradable OR bioabsorbable OR bioresorbable OR degradable OR absorbable OR resorbable) AND (cervical OR thoracic OR lumbar OR thoraco-lumbar OR fusion) AND (plate OR plating OR plates) | 54 | 43 | 11 |
| Science Direct | See above | 7 | 7 | 0 |
| Embase | inion:df or inion:dn | 55 | 54 | 1 |
| Cochrane Central Register of Controlled Trials (CENTRAL) | inion (All text) | 28 | 28 | 0 |
| Clinical trials.gov | inion | 16 | 16 | 0 |
| WHO ICTRP | spine OR spinal / “condition” AND spinal graft containment or S-1 or s 1 or s-2 or s2 / “intervention” | 9 | 9 | 0 |
| Citations referenced in the searched literature | - | 0 | 0 | 0 |
| Articles identified in previous clinical evaluations | - | 1 | 0 | 1 |

Adverse event database search

As part of the clinical evaluation, a search in FDA Manufacturer and User Facility Device Experience (MAUDE) and Total Product Life Cycle (TPLC) databases was performed. The search was conducted to retrieve information of reported adverse events related to the use of Inion Spinal Graft Containment System™. Data was searched from 2005 onwards when the first Inion S-1 products were launched to the market. The search results are copied as unedited in Appendix 8.

As a result, 5 separate reported adverse events related to the use of Inion S-1™ were found.  No adverse events related to the use of Inion S-2TM were found. Results are discussed in Chapter 5.3 PMS data.

Summary and appraisal of the clinical data

The following chapters 4.1-4.3, 5.1 and 5.2 describe the clinical data appraisal and analysis process in more detail.

Criteria for suitability

The publications were screened based on their relevance to the intended use of Inion Spinal Graft Containment System™ in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts. Criteria for suitability are presented in Table 10.

Table 10. Suitability criteria

| Suitability criteria | Description | Grading system |
| --- | --- | --- |
| Appropriate device | Were the data generated from the device in question? | Actual device |
|  |  | Equivalent device |
|  |  | Other device |
|  |  | Insufficient information |
| Appropriate device application | Was the device used for the same intended use? | Same use |
|  |  | Minor deviation |
|  |  | Major deviation |
|  |  | Insufficient information |
| Appropriate patient group | Where the data generated from a patient group that is representative of the intended treatment population and clinical condition? | Applicable |
|  |  | Limited |
|  |  | Different population |
|  |  | Insufficient information |
| Acceptable report/data collation | Do the reports or collations of data contain sufficient information to be able to undertake a rational and objective assessment? | High quality |
|  |  | Minor deficiencies |
|  |  | Insufficient information |

Criteria for data contribution

The criteria for data contribution are presented in Table 11.

Table 11. Criteria for data contribution

| Data Contribution criteria | Description | Grading System |
| --- | --- | --- |
| Data source type | Was the design of the study appropriate? | yes |
|  |  | no |
|  |  | insufficient information |
| Outcome measures | Does the outcome measures reported reflect the intended performance of the device? | yes |
|  |  | no |
|  |  | insufficient information |
| Follow up | Is the duration of follow up long enough to assess whether duration of treatment effects and identify complications? | yes |
|  |  | no |
|  |  | insufficient information |
| Statistical significance | Has a statistical analysis of the data been provided and is it appropriate? | yes |
|  |  | no |
|  |  | insufficient information |
| Clinical significance | Was magnitude of the treatment affect observed clinically significant? | yes |
|  |  | no |
|  |  | insufficient information |

Selection criteria

The final selection was performed as described in Table 12.

Table 12. Selection criteria

| Selection |  |  |
| --- | --- | --- |
| Selection | Description | Abbr. |
| Excluded before assessment | Irrelevant article considering the inclusion criteria, can be excluded based on the title/abstract | EBA |
| Excluded after assessment | Irrelevant article considering the inclusion criteria, can be excluded based on the content of article | EAA |
| Included | Relevant article, included in evaluation/conclusion | INC |

After the data selection, 13 publications were chosen for appraisal. The screening, appraisal processes and results are shown in Appendix 7. After the appraisal, 6 articles were selected for the final analysis. A tabulation of the clinical data used in the evaluation is presented in Table 13. The data is categorized according to whether the data address the performance or the safety, or both, of the device in question. The patient quantities are presented in Chapter 5 when analyzing performance and safety clinical data.

Table 13. Articles included for evaluation

| Article search No. | Authors/ Title/Citations | Does the data address performance | Does the data address safety |
| --- | --- | --- | --- |
| Pub 2 | Rodrigo, et al. Long-term follow-up of anterior cervical discectomy and fusion with bioabsorbable plates and screws. Clin Neurol Neurosurg. 2015 Sep;136:116-21. | Yes | Yes |
| Pub 13 | Truumees E. Commentary: Incorporating new technology into spine care: the case of resorbable cervical plates. Spine J. 2011 Nov;11(11):1009-14. | Yes | Yes |
| Pub 14 | Lebl DR, et al. Bioabsorbable anterior cervical plate fixation for single-level degenerative disorders: early clinical and radiographic experience. Spine J. 2011 Nov;11(11):1002-8. | Yes | Yes |
| Pub 18 | Nabhan A, et al. Comparison of bioresorbable and titanium plates in cervical spinal fusion: early radiologic and clinical results. J Spinal Disord Tech. 2009 May;22(3):155-61. | Yes | Yes |
| Pub 21 | Franco A et al. Use of resorbable implants for symptomatic cervical spondylosis: experience on 16 consecutive patients. J Neurosurg Sci. 2007 Dec;51(4):169-75. | Yes | Yes |
| Cit 1 | Arianne P et al, 2013: Effect of various kinds of cervical spinal surgery on clinical outcomes: A systematic review and meta-analysis | Review | Review |

Data analysis

Overall, the selected clinical data consisted of 6 publications, including 1 review, 1 commentary and 4 clinical studies in which at least 66 patients were followed up with average follow–up time of 2.2 years. The data consisted of performance and/or safety data about actual device: Inion Spinal Graft Containment System™ implants and instruments. Publications which demonstrated performance are included in chapter 5.1 Performance, while publications which contain data suitable for safety analysis are considered in 5.2 Safety.

Table 14. A systematic review

| Authors/ Title | Details | About Inion Spinal Graft Containment SystemTM |
| --- | --- | --- |
| Arianne P et al, 2013: Effect of various kinds of cervical spinal surgery on clinical outcomes: A systematic review and meta-analysis | Article Type: Systematic review Objective: to evaluate the benefits and harms of surgical techniques in comparison with each other in patients with cervical disorders based on patient-reported outcomes such as pain, function, recovery, and reoperation. Main Results: | Reference studied Inion Spinal Graft Containment System™: Pub 18. |

There was one systematic review included in the data analysis. The review was conducted by Arianne P et al. (2013), aimed to assess the benefits and harms of cervical spinal surgery in patients with cervical disorders. The study by Nabhan et al. (2009) was one of the 39 studies which included in this review, and meta-analysis comparing different surgical interventions in patients with disorders of the cervical spine. Based on their findings the authors concluded that differences in benefits and harms between the various surgical techniques are small and therefore the surgeon, patient, and health care provider can make the choice of any surgical technique based on experience, preferences, or costs.

Post-Market Surveillance (PMS) data analysis from FDA MAUDE database and Inion’s complaint handling records is presented in chapter 5.3.

Performance

The Essential Requirement 3 of MDD pertains to the clinical performance of Inion Spinal Graft Containment System™ (implants and instruments) which is looked at in this evaluation. The articles related to performance are listed below in Table 15.

Table 15. Performance related articles

| Authors/ Title | Details (Study type /product /indications /Objective) | Patient qty. | Follow-up time | Performance |
| --- | --- | --- | --- | --- |
| Franco A et al. Use of resorbable implants for symptomatic cervical spondylosis: experience on 16 consecutive patients. | Study type: Retrospective study Product: Inion S-1 Indications: Anterior cervical Objective: to evaluate the results of a consecutive series of 16 patients affected by degenerative cervical spondylosis and operated on by anterior cervical discectomy and fusion (ACFD) by means of anterior bioresorbable plates and screws. | 16 | 12-25 months | Fusion rate with biodegradable plate: 90% vs. without plate: 88.9% |
| Nabhan A, et al. Comparison of bioresorbable and titanium plates in cervical spinal fusion: early radiologic and clinical results. | Study type: a randomized, controlled, prospective study Product: Inion S-1 Indications: Anterior cervical Objective: to compare clinical results, segmental motility, magnetic resonance imaging (MRI) compatibility, and change of the bone density of a cervical spine segment that was treated with either bioresorbable or titanium plates in single level. | 19 | 6 months | - Bioresorbable plate shows comparable fusion characteristics to titanium plate, fusion rate was 94.6%; - No removal surgery needed for S-1 products; - Lucency and radiolucency of bioresorbable plates benefits for surgical procedure. |
| Lebl DR, et al. Bioabsorbable anterior cervical plate fixation for single-level degenerative disorders: early clinical and radiographic experience. | Study type: Retrospective comparative cohort study Product: Inion S-1 Indications: Anterior cervical Objective: To compare the clinical and radiographic outcomes of patients undergoing ACDF for single-level degenerative disorders with a bACP versus a conventional metal ACP (mACP). | 14 | 11.4±2.6 months | Patients with the resorbable plate had higher pseudathrosis rates (29% vs. 0%); lesser ability to maintain segmental lordosis (21% vs. 53%); and poorer clinical outcomes (50% good/excellent vs. 93%). |
| Truumees E. Commentary: Incorporating new technology into spine care: the case of resorbable cervical plates. | Study type: commentary Product: Inion S-1 Indications: Anterior cervical Objective: Commentary for article-Lebl DR, et al, 2011 | -- | -- | -- |
| Rodrigo, et al. Long-term follow-up of anterior cervical discectomy and fusion with bioabsorbable plates and screws. | Study type: a retrospective descriptive study Product: Inion S-1 Indications: Anterior cervical Objective: to evaluate the long-term clinical results of the ACDF implants made out of bioabsorbable materials. | 17 | 5-7 | Complete fusion was achieved in 100% of the cases. |

The first clinical study regarding Inion Spinal Graft Containment System™ was conducted by Franco et al. in 2007, to evaluate the results of 16 patients affected by degenerative cervical spondylosis and operated on by anterior cervical discectomy and fusion (ACFD) by means of anterior bioresorbable plates and screws. Further, the authors compared the results in these patients with 13 patients with same indication in whom arthrodesis was obtained by using cages without plates. 8 males and 8 females aging from 37 to 69 years were included in Group A, and 9 males, 4 females aging from 50 to 77 years were in the Group B. Group A was treated with dense cancellous allografts (Graftech; Osteotech; Inc.) and anterior resorbable plates and screws plates (INION OY) Surgical procedure was carried out at one level in 11 patients and at two levels in the remaining 5 cases. During the same period, other 13 patients (Group B) were operated by using intervertebral cages Solis (Stryker Spine), at one level in 10 patients and at two levels in 3 patients. The follow-up ranged from 12 to 25 months.

The fusion rates and the vertebral alignment were evaluated. For the vertebral body and disc space height (DSH), in the Group A, the mean values were: 37.5 mm (range 30-51 mm) soon after surgery and 36.9 mm (range 27-49 mm) at the most recent follow-up; in the Group B patients, the same measurement revealed mean values of 34.3 mm in the early postoperative period (range 21-49 mm) and of 25.7 mm (range 20-44 mm) at the most recent follow-up. Finally, the data relative to the vertebral alignment showed mean values of 2.75 soon after surgery and of 5.31 in the most recent follow-up in Group A. Such data proved the efficacy of these systems to ensure a good degree of fusion while maintaining vertebral alignment and ultimately stability in the operated segment. In the patients observed in this study the authors inserted a dense cancellous allograft stabilized by Inion cervical plate and screws gaining a fusion rate of 90%, while in the patients treated with Solis cage without plates, a fusion rate of 88.9% was achieved. With regard to the fusion rate is had to be noted that resorbable plates showed a modulus of elasticity closer to normal bone, discussed by the authors. The presurgical values of VAS decreased in the late postsurgical period whereas the JOA increased in both group of patients so revealing a meaningful improvement of the symptoms and neurological signs. It was concluded that the resorbable plate system ensures a stabilization of the operated segment long enough to ensure a good fusion before being degraded. Finally the ability of the resorbable plate of not interfering with the neuroradiological postoperative examinations represented a sure advantage over the metal plates.

In study by Nabhan et al. (2009) the Inion Spinal Graft Containment System™ (S-1™ System) was successfully used in anterior cervical discectomy and fusion (ACDF). A prospective, randomized, controlled clinical trial was performed with 37 patients in order to compare bioresorbable (n=19) and titanium (n=18) plate fixation in single-level ACDF. Follow-up used a visual analog scale (VAS) with regard to brachial pain and Neck Disability Index (NDI) for neck pain. Success of the procedure was evaluated by radioisometry (immediately postoperatively, 6 weeks, 3 and 6 months), MRI of cervical spine (immediately postoperatively, 3 and 6 months) and computed tomography. Satisfactory clinical results were achieved in both groups without differences in fusion rates or speeds.

In retrospective study by Lebl et al. (2011) the clinical and radiographic outcomes of patients undergoing ACDF with structural allograft for single-level degenerative disorders with Inion S-1 System was compared to patients treated with a conventional metallic plates and titanium mesh cage filled with morzeliced allograft bone. In both groups patients were immobilized postoperatively in a rigid cervical orthosis for 2 weeks. 14 patients underwent ACDF with Inion S-1 and 15 with a metallic plate. The authors retrospectively reviewed the results of a consecutive series of patients undergoing ACDF for symptomatic single-level disc herniation or spondylotic neural compression over a 3-year period. Operative notes, clinical charts, and radiographs were analyzed. Radiographic outcomes were assessed for intersegmental alignment, graft subsidence, fusion rate, prevertebral soft-tissue shadow, and graft containment. Clinical outcome was evaluated by Odom’s criteria. Radiographic outcomes at the most recent follow-up demonstrated pseudarthrosis in 4 of 14 patients (29%) in the S-1 group and 0 of 15 patients in the metallic group. Comparing preoperative and final radiographs, cervical lordosis was maintained at the operative segment in only 3 of 14 S-1 patients (21%) compared with 8 of 15 patients (53%) in the metallic group. The mean Cobb angle was 2.4°±1.9° lordosis in the S-1 group and -2.7°±2.5° kyphosis in the metallic group (p=0.12). In the metallic group, 14 of 15 patients had good or excellent results. In the S-1 group, only 7 of 14 patients had good or excellent results. The details were following:

- Odom's criteria:

  - 3 poor (21%) (one smoker)

  - 4 fair (29%)

  - 3 good (21%)

  - 4 excellent (29%)

In this study bioabsorbable ACP fixation, i.e., Inion S-1 fixation, was associated with a high rate of graft extrusion and early loss of intersegmental cervical alignment. Inferior clinical outcomes were observed in patients in the S-1 group compared with the metallic group. Based on their findings, authors stated and concluded that continued use of the Inion S-1 System cannot be recommended.

Truumees (2011) published his commentary article considering usage of resorbable cervical plates in spine surgery. In this review-type article no new data considering the Inion S-1 system was presented but study by Lebl et al. (2011) mentioned above was referenced. Author pointed out the importance of appropriate use of resorbable cervical plates and discussed about obvious risks related to their off-label use and limitations in their mechanical properties in scope of load-bearing applications. Truumees also criticizes FDA’s 510(k) clearance process in case of resorbable spinal devices and wonders the limited amount of clinical data prior their approval for commercial use. On the other hand, Truumees also pointed out that resorbable plates are proposed to compensate for the potential downsides of metal anterior cervical plates, such as graft shielding from axial loads, increasing pseudoarthrosis and the risk of adjacent segment degeneration.

Goal of study by Rodrigo et al. (2015) was to investigate long term effectiveness of Inion S-1™ System for the anterior cervical discectomy and fusion (ACDF) in treatment of cervical spinal degenerative disease. Total of 17 patients (ten women and seven men, age range 30-58 years with mean of 45 years) underwent ACDF using Inion S-1™ resorbable plates and screws. Total of 13 one level and four dual level ACDF were performed. The most common level operated was C5-C6. To evaluate results and outcomes medical records, pre- and post-surgical symptoms, previous trauma, smokers, complications and prevertebral space in pre- and post-intervention were analysed. Patients remained hospital 2-8 days post-surgery (mean 3.5 days, 76% 3 days or less). Follow-up up was conducted up to 5-7 years (2006-2013). Based on their results authors concluded that the study provides proof for easy handling of Inion S-1™ Plates and good tolerance for their material, i.e., biodegradable copolymer of 80:20 poly (L-lactide-co-D,L-lactide). Author stated that such materials offer a feasible alternative to classical ACDF as the body itself integrates conventional titanium plates, but titanium implants remain in the body while the resorbable material is degraded and converted into bone.

Among the articles analyzed above, only one clinical study conducted by Lebl et al. (2011) showed inferior clinical outcomes in the S-1 group compared with the metallic group. These Inion S-1 products are intended to be used with additional rigid fixation (posterior interspinous wiring). It is noted though, that a similar Inion S-1 system is available outside the US, intended to be used in connection with a postoperative collar. In Lebl study additional posterior rigid fixation was not used; and postoperative collar was used for only for 2 weeks. It is estimated that this is improper use for even those indications, as the more general time period for postoperative collar after ACDF is approximately 3-6 weeks but is always determined by the surgeon for each case.

It is further noted that complications occur also when using devices made of metal, requiring re-operation and implant removal. With reference to Lowery et al. (1998) with up to 35% failure rate for metal plates in anterior cervical fusion.  Lowery et al. conclude that hardware failure associated with anterior cervical surgery does not pose a significant risk to the patients. In the case of nonabsorbable implants, careful and long-term follow-up can ensure that hardware has not progressed and can confirm that late failure has not occurred.  It is also clearly stated in the IFU that these devices (unlike metal plates) are not intended for full load bearing: "WARNINGS:• The INION S-1™ BIODEGRADABLE ANTERIOR CERVICAL FUSION SYSTEM devices provide fixation and are not intended to replace healthy tissues or withstand the stress of full load bearing."

Nabhan et al. 2009 study shows that when used correctly in accordance with the current instructions for use as detailed above, the anterior plate fixation using Inion S-1 plate has the same fusion progress and stability as titanium, with no biocompatibility issues, and with several benefits in comparison to metal such as: ability to contour the plate to patient anatomy, no imaging artifact, and ability to see the graft and vertebrae in postoperative imaging, implant removal is unnecessary in future surgeries (drill-through and degradation), material's modulus of elasticity allowing load sharing with the graft.

These results are supported by Rodrigo et al. (2015) study, who investigated long term effectiveness of Inion S-1™ System for the anterior cervical discectomy and fusion (ACDF) in treatment of cervical spinal degenerative disease. Total of 17 patients (ten women and seven men, age range 30-58 years with mean of 45 years) underwent ACDF using Inion S-1™ resorbable plates and screws. Total of 13 one level and four dual level ACDF were performed. Follow-up period was 5-7 years. None of the patients had associated complications such as edema in adjacent tissues or long-term infections. All segments have documented fusion. No patient had to be reoperated due to failure or migration of the implants used in surgery. The authors conclude several benefits of the biodegradable materials in this application. The bioabsorbable implants allow gradual distribution and transfer of loads to the surrounding tissues minimizing the riskof stress shielding during the early stages of fusion (allowing correct fusion between the graft and fixed bodies) prior to the process of degradation of the plates and screws. Flexible and less rigid bioabsorbable implants provide the advantage of allowing a greater transfer of load to the healing bone over time during the degradation of the implants, potentially minimizing the risk of degeneration at adjacent levels. Other potential advantages of resorbable implants in spinal surgery include a more physiological modulus of elasticity in relation to native vertebral bone allowing more optimal load-sharing, as compared with rigid metal devices, and the elimination of long-term residual metal implants and their potential complications (migration or dislodgement) and the need for secondary hardware removal (e.g., in case of later adjacent level procedures). Biodegradable implants allow optimal postoperative radiographic evaluation because of their radiolucency and the absence of artifacts associated with similar metallic devices exposed to advanced imaging modalities. However, the postoperative position of implanted biodegradable implants can be evaluated radiographically if the implants contain radiopaque markers.

No new performance risks have been discovered during this evaluation. Moreover, the risk analysis, verification and validation processes and post-marketing surveillance ensure that Inion Spinal Graft Containment System™ complies with performance requirements.

Safety

The Essential Requirement 1 of MDD pertains to the clinical safety of Inion Spinal Graft Containment System™ implants and instruments which is looked at in this evaluation.  The articles related to safety listed below in Table 16.

Table 16. Safety related articles

| Authors/ Title | Details (Study type /product /indications /Objective) | Patient qty. | Follow-up time | Safety |
| --- | --- | --- | --- | --- |
| Franco A et al. Use of resorbable implants for symptomatic cervical spondylosis: experience on 16 consecutive patients. | Study type: Retrospective study Product: Inion S-1 Indications: Anterior cervical Objective: to evaluate the results of a consecutive series of 16 patients affected by degenerative cervical spondylosis and operated on by anterior cervical discectomy and fusion (ACFD) by means of anterior bioresorbable plates and screws. | 16 | 12-25 months | No incidence of postoperative hematoma or infection. |
| Nabhan A, et al. Comparison of bioresorbable and titanium plates in cervical spinal fusion: early radiologic and clinical results. | Study type: a randomized, controlled, prospective study Product: Inion S-1 Indications: Anterior cervical Objective: to compare clinical results, segmental motility, magnetic resonance imaging (MRI) compatibility, and change of the bone density of a cervical spine segment that was treated with either bioresorbable or titanium plates in single level. | 19 | 6 months | MRI revealed no signs of soft tissue inflammation or swelling in both groups 6 months after surgery. |
| Lebl DR, et al. Bioabsorbable anterior cervical plate fixation for single-level degenerative disorders: early clinical and radiographic experience. | Study type: Retrospective comparative cohort study Product: Inion S-1 Indications: Anterior cervical Objective: To compare the clinical and radiographic outcomes of patients undergoing ACDF for single-level degenerative disorders with a bACP versus a conventional metal ACP (mACP). | 14 | 11.4±2.6 months | - There were two patients with dysphagia in the bioresorbable group and one patient with dysphagia in the metal group. - Patients with the resorbable plate had higher graft extrusion in 75% of pseudarthroses; |
| Truumees E. Commentary: Incorporating new technology into spine care: the case of resorbable cervical plates. | Study type: commentary Product: Inion S-1 Indications: Anterior cervical Objective: Commentary for article-Lebl DR, et al, 2011 | -- | -- | -- |
| Rodrigo, et al. Long-term follow-up of anterior cervical discectomy and fusion with bioabsorbable plates and screws. | Study type: a retrospective descriptive study Product: Inion S-1 Indications: Anterior cervical Objective: to evaluate the long-term clinical results of the ACDF implants made out of bioabsorbablematerials. | 17 | 5-7 years | - No patient had to be reoperated because of extrusion, breach or implant breakage. - No postoperative hematoma, wound infection, osteomyelitis or recurrent laryngeal paresis. - No patient reported fever, chills or itching. - No material-related inflamma-tion was found in our study. |

Franco et al. (2007) conducted a following testing protocol: Cervical fusions with biodegradable Inion plates (16 patients, Group A) vs. allograft only (13 patients, Group B) to treat degenerative cervical spondylosis. The method and follow-up are described in 5.1. During the following up time, the postoperative complications including dysphonia, dysphagia, neck and brachial pain and the infection rate was paid particular attention. While for the result, no symptoms of local pain and/or swelling were found during the healing. Moreover, degradation of the implants was not observed to cause harmful inflammatory response. The reaction of the soft tissues was evaluated by measuring the values of soft tissue space, which were 24.1 mm soon after surgery (range 13-39 mm) and 17.8 mm at the most recent follow-up (range 8-35 mm) in Group A cases, and 22.7 mm (range 13-25 mm) soon after surgery and 17.2 mm (range 7-31 mm) at the last follow-up in Group B cases. These decreased values also confirmed that in the cases analyzed, the resorption of bioresorbable implant was not affected by any adverse reaction.

In study by Nabhan et al. (2009) the Inion Spinal Graft Containment System™ (S-1™ System) was successfully used in anterior cervical discectomy and fusion (ACDF). A prospective, randomized, controlled clinical trial was performed with 37 patients in order to compare bioresorbable (n=19) and titanium (n=18) plate fixation in single-level ACDF. During the 6 months of following-up, no complications like soft tissue inflammation, swelling and inflammation were observed.

In retrospective study by Lebl et al. (2011) the clinical and radiographic outcomes of patients undergoing ACDF with structural allograft for single-level degenerative disorders with Inion S-1 System was compared to patients treated with a conventional metallic plates and titanium mesh cage filled with morzeliced allograft bone. In both groups patients were immobilized postoperatively in a rigid cervical orthosis for 2 weeks. 14 patients underwent ACDF with Inion S-1 and 15 with a metallic plate. The radiographic outcomes at the most recent follow-up demonstrated:

- 3 with nonunion/extruded graft at 4.5 - 7 months, revision surgery

- 1 dysphagia - return to clinic for sterile fluid collection at 1 month follow up

- 1 intermittent dysphagia

- 1 delayed union (smoker)

- 1 nonunion at 11 months

For the conclusion, authors stated and concluded that continued use of the Inion S-1 System cannot be recommended.

Truumees (2011) published his commentary article considering usage of resorbable cervical plates in spine surgery. In this review-type article no new data considering the Inion S-1 system was presented but study by Lebl et al. (2011) mentioned above was referenced. The Trumees and Lebl studies are discussed in more detail above in chapter 5.1. Performance. Concerning safety events in Lebl, there was one case of dysphagia where patient returned to clinic for sterile fluid collection one month postoperatively. It is estimated that this short term tissue reaction is not related to implant degradation, and also other materials which were implanted during this operation may have caused or contributed to the sterile fluid accumulation.

Goal of study by Rodrigo et al. (2015) was to investigate long term effectiveness of Inion S-1™ System for the anterior cervical discectomy and fusion (ACDF) in treatment of cervical spinal degenerative disease. Total of 17 patients (ten women and seven men, age range 30-58 years with mean of 45 years) underwent ACDF using Inion S-1™ resorbable plates and screws. Total of 13 one level and four dual level ACDF were performed. Follow-up up was conducted up to 5-7 years (2006-2013). As results, none of the patients had associated complications such as edema in adjacent tissues or long-term infections. All segments have documented fusion. No patient had to be reoperated due to failure or migration of the implants used in surgery.

It is noted that although there are no published studies identified with Inion S-2 products (graft containment plates and screws for anterior thoraco-lumbar fusion), they are equivalent products with Inion S-1 products (graft containment plates and screws for anterior cervical fusion) and thereby their safety and performance is evidenced by the safety and performance of the Inion S-1 products. Both subgroups are intended to be used in conjunction with posterior rigid fixation. This similarity of the subsystems is further evidenced by the similar or lower occurrence rate of adverse events and complaints, i.e. there are no cases of post-operative problems with Inion S-2, and all events are related to problems during insertion (please see section 5.3 below).

The equivalency of Inion S-1 and Inion S-2 products is listed herewith:

- same indications for use, only for different level of application in the spine

- same material

- same manufacturing method, equipment and facility

- same principle of operation

- similar geometry/design

- similar instrumentation.

No new safety risks have been discovered during this evaluation. In addition, the risk analysis, verification and validation processes ensure that Inion Spinal Graft Containment System™ complies with the safety requirements.

PMS data

Inion S-1TM has been in clinical use for 12 years, and more than 1500 operations have been conducted with these products. Inion S-2TM was introduced in 2007, and the use of these implants has started to increase only recently, during 2015-2016. Approximately 200 operations have been conducted with the 2-hole plates which were introduced as a line extension. The post-market clinical follow-up data is analyzed through sales vs. complaints data, customer feedback investigations, adverse event reports and clinical literature evaluation.

5.3.1 Sales Information

The sales numbers, reported incidents, and occurrence rates of (any type of) patient harm are presented in Tables 17 and 18 below.

Table 17. Inion S-1: Sales and risk occurrence rate, long term

| Inion S-1 | 2005-2016 |
| --- | --- |
| Total units sold (screws and plates) | 6727 |
| Estimated No. of operations | 1543 |
| Occurrence of any patient harm | 24 |
| Reported incidents | 4 |
| Occurrence rate: | 1.6% |

Table 18. Inion S-2: Sales and risk occurrence rate, long term

| Inion S-2 | 2007-2016 |
| --- | --- |
| Total units sold (screws and plates) | 557 |
| Estimated No. of operations | 214 |
| Occurrence of any patient harm | 1 |
| Reported incidents | 0 |
| Occurrence rate: | 0.5% |

The long term ratio of patient related complaints/sales has been constant and continuously low: 1.6% for Inion S-1 and 0.5% for Inion S-2, since the products were launched to the market. There is no clinically relevant change in the trend over time.

5.3.2 Adverse events and recalls

There have been no recalls or field safety corrective actions concerning the Inion Spinal Graft Containment System products. A search was made in Inion QMS CAPA records as well as FDA Total Product Life Cycle database (TPLC). The recalls were searched by using keywords “Inion S-1”, “Inion S-2” and FDA product code “OJB”.

Reported adverse events for the Inion Spinal Graft Containment System were searched from FDA Manufacturer and User Facility Device Experience (FDA Maude) database from 2005 onwards. Six reports were identified related to Inion S-1, of which number 1 and 2 represent the same event, reported by both Inion and surgeon. Number 5 is reported by the surgeon while Inion has evaluated that this is not a reportable adverse event.  No reported events were identified related to Inion S-2 (Table 19). These reports match with the records of Inion QMS Complaint records i.e. all issues have been formally evaluated during Inion complaint handling process.

Table 19. Adverse events from the FDA Maude and TPLC databases (2005-2016).

| No. | Report received by FDA | Device | Event Type | Event Description summary | Hazard/Risks |
| --- | --- | --- | --- | --- | --- |
| 1 | 2009 | INION S-1 GRAFT CONTAINMENT SYSTEM PLATE (Inion CFI No. CFI-09-002) | Injury | About 8 months after surgery, a patient appeared to have some resorption of bone related to bioresorbable plate and screws. Patient did not require re-operation but is only now starting to fuse more solidly. A second patient had a stable fusion but a cystic mass at the site of the plate that had to be removed due to swallowing difficulties. Possibly a third case with evidence on mri of both inadequate bone growth and cystic mass at the site of the plate. | tissue reaction |
| 2 | 2009 | INION S-1 GRAFT CONTAINMENT SYSTEM PLATE (Inion CFI No. CFI-09-002) | Injury | Acdf, c5-6. Solid fusion, but swallowing difficulties, abnormal cystic mass (same shape and size as the plate) and liquid formation 1. 5 years post-operatively. Re-operation performed for cyst removal (inflammatory tissue). Patient is fine. | tissue reaction |
| 3 | 2011 | INION S-1 BIODEGR. GRAFT CONTAINMENT PLATE (Inion CFI No. CFI-11-015 USA) | Injury | Patient had a c5-c7 inion plate for an acdf. The patient developed loss of her lordotic curvature after the surgery on f/u films: reversal of lordosis at c5-7. Partial collapse of c4/5 also. In the revision surgery the inion plate and screws were removed. The screw heads were found loose which according to surgeon may or may not have occurred during dissecting fibrous tissue off the plate. The plate cracked into two during removal. Surgeon conclusion is loss of bone structure. The patient had been in motor vehicle accident prior to the operation. | Fixation failure, non-union, malunion |
| 4 | 2011 | INION S-1 BIODEGR. GRAFT CONTAINMENT PLATE (Inion CFI No. CFI-11-017) | Injury | After a c5/6 acdf surgery, debridement and antibiotics 2 weeks post operative. Revision surgery of soft tissue abscess and implant removal 3 months post op. Observations: scar tissue in neck, no infection. Plate intact. Collapse of c5 onto cage seen in (b)(6) 2011 film (before the revision). Patient condition is reported good. | Fixation failure, non-union, malunion |
| 5 | 2014 | INION S-1 BIODEGR. GRAFT CONTAINMENT PLATE (Inion CFI No. CFI-14-010) | malfunction | Patient in surgery for anterior cervical discectomy with fusion. During surgery, 3 inion cervical plates cracked while surgeon was posting them to the patient's cervical vertebrae. It appears the correct temperature of the water may not have been met. | Problems during insertion |
| 6 | 2015 | INION S-1 BIODEGR. GRAFT CONTAINMENT PLATE (Inion CFI No. CFI-15-015) | Injury | In the 8 weeks post op x-rays, surgeon observed graft subsidence in connection with acdf procedure, where inion s-1 plate and screws were used as graft containment device. Screws were observed to be broken below the plate from the upper end of the plate. In 2 weeks post-op pictures the fixation had been observed as fine. The patient was wearing a hard collar for 3 weeks after the operation. The patient has not had any symptoms, but a revision surgery was conducted to avoid any complications later. | Fixation failure, non-union, malunion |

The reported adverse events represent the recognized risks in the risk management report of Inion Spinal Graft Containment system. The occurrence of these hazards is presented in Table 20 below.

Table 20. Risk occurrence per hazard type, reported adverse events in FDA Maude

| Hazard/Risk | Total | Percentage of all events |
| --- | --- | --- |
| Problems during insertion/implantation | 1 | 20% |
| Implant has toxic or injurious effect | 0 | 0% |
| Infection | 0 | 0% |
| Tissue reaction | 1 | 20% |
| Fixation failure, non-union, malunion | 3 | 60% |
| Undesirable outcome due to improper use | 0 | 0% |
| Total | 5 | 100% |

5.3.3 Customer complaints and feedback

The customer complaints and feedback received and investigated by the manufacturer during 2007 - 2016 are categorized and discussed in this chapter.

There have been altogether 27 separate customer feedback investigations during the time period. These include the evaluation of patient cases from the Lebl 2011 study discussed in the clinical data evaluation above. Based on the customer feedback data, the hazard analysis is shown in Tables 21 and 22. Unlike Tables 17 and 18, these tables include also complaints which caused no actual harm to the patient except a slightly longer operation time which is clinically irrelevant (such as implant breakage during insertion, other insertion problems, and instrument related problems).

Table 21. Inion S-1: Hazards analysis of customer feedback evaluations

| Hazard/Risk | Events | Occurrence rate (events/total cases) |
| --- | --- | --- |
| Fixation breakage, nonunion, malunion | 12 | 0,8% |
| Infection, tissue reaction / short term | 1 | 0,1% |
| Tissue reaction / long term | 4 | 0,3% |
| Mechanical irritation | 0 | 0,0% |
| Graft subsidence | 5 | 0,3% |
| Implant breakage during insertion | 9 | 0,6% |
| Other insertion problems and instrument issues | 3 | 0,2% |
| Total | 34 | 1452 |

Table 22. Inion S-2: Hazards analysis of customer feedback evaluations

| Hazard/Risk | Events | Occurrence rate (events/total cases) |
| --- | --- | --- |
| Fixation breakage, nonunion, malunion | 0 | 0,0% |
| Infection, tissue reaction / short term | 1 | 0,5% |
| Tissue reaction / long term | 0 | 0,0% |
| Mechanical irritation | 0 | 0,0% |
| Graft subsidence | 0 | 0,0% |
| Implant breakage during insertion | 1 | 0,5% |
| Other insertion problems and instrument issues | 5 | 2,3% |
| Total | 7 | 214 |

Customer complaints are usually received from users (surgeons) or distributors. Customer feedback evaluation shows that most of the product problems were related to non-optimal usage technique or implant selection, which refers often to the fact that there is a learning curve when surgical operations are conducted with biodegradable implants instead of titanium implants. Such events tend to decrease as the surgeons gain more experience on the usage of these materials and insertion techniques.

Risk of adverse reactions is always present in surgical operations when foreign material is implanted regardless of implant material, size, and operation site.

5.3.4 Clinical Application in Literatures

Related safety reports are described in 5.2.

Inion Spinal Graft Containment System has been on the market for 10 years, and it has achieved good results in clinical use. The manufacturing process, materials and insertion accessories feature good biocompatibility; and the structural design of products shows safety and reliability, to meet the needs of clinical use. No new risks have been found during the clinical use.

5.3.5 Conclusion on PMS data

Based on Inion’s PMS data of the clinical use of Inion S-1TM, the complication/operations rate in 10-year follow up (2005-2016) is 1.6%, which is considered low and acceptable. Inion’s PMS data of the clinical use of Inion S-2TM shows that the complication/operations rate in 9-year follow up (2007-2016) is 0,5% which is considered low and acceptable. The number of events has been reducing over time, as surgeons have gained more experience and knowledge of the insertion techniques, implant selection and patient selection.

All the reported adverse events have been anticipated residual risks. Inion has provided risk control measures to the anticipated residual risks in the Risk Management Report - Inion Spinal Graft Containment System and Instruction for Use. Analysis of the clinical literature, adverse events and customer complaints data show that the Inion Spinal Graft Containment System meets the requirements of clinical evaluation and verification. The risks identified in Risk management document are shown to be in control by the clinical evaluation. The PMS data demonstrates that the Inion Spinal Graft Containment System is safe and effective in its intended use.

Requirement on acceptable benefit/risk profile

The Essential Requirement 1 of MDD pertains to the acceptable benefit/risk profile of Inion Spinal Graft Containment System™ which is looked at in this evaluation.

Based on the pertinent clinical data from literature and post-marketing data of Inion Spinal Graft Containment System™ products, it can be concluded that no significant health or safety risk was discovered during this evaluation.  The overall complication rates and risks related to these products remains continuously low and acceptable, and there are no clinically relevant change is detected in the trends over time. Therefore reviewed data provides evidence that complication rates and risks related to Inion Spinal Graft Containment System™ can be justifiable stated to be low and acceptable, and that the Inion Spinal Graft Containment System™ products are and can be expected to be clinically safe and effective for their intended use when used according to their product-specific instructions for use.

In addition, the risk analysis, verification and validation processes ensure that Inion Spinal Graft Containment System™ benefit/risk profile complies with the clinical safety and performance requirements.

Requirement on acceptability of side effects

The Essential Requirement 6 of MDD pertains to the acceptability of side effects of Inion Spinal Graft Containment System™ which is looked at in this evaluation.

Based on the evaluated clinical data, no new side effects have been discovered during this evaluation. Anticipated residual risks may occur, but the number is low. Inion has provided risk control measures to the anticipated residual risks in the Risk Management Report - Inion Spinal Graft Containment System™ and Instruction for Use.

Also, the risk analysis and biological safety evaluation ensure that Inion Spinal Graft Containment System™ complies with the safety and performance requirements.

Product documentation and instructions for use

The manufacturer’s product documentation and instructions for use are well in line and consistent with the clinical data of the device use and indications. The texts cover all the hazards and other clinically relevant information that may impact the use of the device.

Conclusions

This clinical evaluation complied with the MDD Essential Requirements 1, 3 and 6.

As a summary of findings about Inion Spinal Graft Containment System™’s clinical safety and performance, the following conclusions are made:

The residual risks identified in the risk management documentation have been addressed by this clinical evaluation. The published literature concerning the actual devices in clinical use provides sufficient data to conclude that Inion Spinal Graft Containment System™ with accessory instruments are clinically safe and efficient for their intended use when used according to their product-specific Instructions for Use.

Evidence demonstrates the material safety of resorbable polylactic acid copolymer P(L/DL)LA. The mechanical strength and degradation characteristics of these implants are sufficient and appropriate to maintain the relative position of weak bony tissue such as allografts and autografts, which have been tested and verified.

Possible adverse events include tissue reaction, non-union or malunion. Although tissue reaction is a known risk (adverse event, anticipated risk) which always exists with all biodegradable implants, these reactions typically occur first 1-2 years postoperatively and do not affect bone healing negatively or cause permanent damage or impairment. Tissue reactions are possible when the device degrades, especially when a plate is implanted under a thin soft tissue layer. However, the occurrence rate has been reported to be very low with polylactic acid based implants. Risk control measures have been taken to the anticipated residual risks during the risk management process.

The products represent the current state of the art in spinal fusion procedures as a means to maintain the relative position of weak bony tissue such as allografts or autografts. It can be also justifiable stated that the clinical product-specific benefits can be considered to clearly overweigh the identified product-specific risks.

Next clinical evaluation

As extensively outlined above, the use of Inion Spinal Graft Containment System™ is well-established and the safety profile is well-known without significant risks. Safety and performance of this product has been examined and documented in many clinical studies. Moreover, extensive experience in clinical practice and post-marketing data support the performance and safety profile of Inion Spinal Graft Containment System™ implants and instruments in the claimed indications. Unknown adverse events caused by the Inion Spinal Graft Containment System™ implants and instruments are not expected to occur. Accordingly, if the manufacturer does not receive any new information from Post-Market Surveillance (PMS) that has the potential to change the current evaluation, the clinical evaluation has to be actively updated every 3 years.

Dates and signatures

The evaluators declare that, to the best of their knowledge, sufficient effort was made to identify and evaluate relevant literature which allowed a clinical evaluation. Based upon the selected information it was possible to critically evaluate the clinical performance, risks and benefits of the device in question. Potentially negative contributions were not excluded from the evaluation. The evaluators agree with the contents of this report.

NAME					SIGNATURE				DATE

Evaluators

______________________________________________

Janne Nurmi, Clinical Representative (DVM, PhD, Adjunct Professor)

Sales Director, Inion Oy, Tampere, Finland

______________________________________________

Qinqin Wang, Regulatory Affairs Specialist & Material Researcher,

Naton Institute of Medical Technology, Beijing, China

______________________________________________

Qiyun Zhou, Regulatory Affairs Specialist,

Naton Institute of Medical Technology, Beijing, China

Final release by the manufacturer is demonstrated by approving and publishing this document in Inion QMS as current document (MULN-7JJCQ2) revision 6.

Qualification of the responsible evaluator

The manufacturer acknowledges that the evaluator of this clinical evaluation has the required qualifications and professional experience to conduct the clinical evaluation of Inion Spinal Graft Containment System™ implants and instruments. The relevant documents on qualification and experience are held by the manufacturer.

Declaration of interests

The Evaluators and Authors have made all effort to provide an objective evaluation in accordance with the clinical evaluation plan. Evaluators and authors are employees of Inion Oy and of the group company Naton Medical Group and its subsidiaries.

References

Appendices

Appendix 1 - <Product Name> - PubMed raw data

Appendix 2 - <Product Name> - ScienceDirect raw data

Appendix 3 - <Product Name> - Embase raw data

Appendix 4 - <Product Name> - Cochrane Reviews raw data

Appendix 5 - <Product Name> - Cochrane Trials raw data

Appendix 6 - <Product Name> - Clinicaltrials.gov and WHO ICTRP raw data

Appendix 7 - <Product Name> – Data search results and data selection and data appraisal

Appendix 8 - <Product Name> – FDA Maude raw data