## Slide 1: Applying the screening panel decision criteria in the context of the CECP
Training of members of Commission expert panels on medical devices and in vitro diagnostic devices (EXPAMED)

## Slide 2
1
2
3
4
5
Background and introduction
Criterion 3: significantly increased incidents
Criterion 1: novelty and resulting impacts
Criterion 2: scientifically valid health concerns
Overall assessment and available tools

## Slide 3
1
2
3
4
5
Background and introduction
Criterion 3: significantly increased incidents
Criterion 1: novelty and resulting impacts
Criterion 2: scientifically valid health concerns
Overall assessment and available tools

## Slide 4
Pre-market space
Post-market space
Post-market clinical follow-up
Post-market surveillance
When do expert panels come into play?
Scientific opinion on CEAR
Expert Panels
Notified Body
Clinical evaluation assessment report (CEAR)
Conformity assessment
Manufacturer
Documentation for conformity assessment
Notified Body
Consider expert panel opinion for finalising conformity assessment
Issue certificate of conformity

## Slide 5
Notified Body
Consider legal requirement
Exemptions
Clinical evaluation consultation procedure (CECP)
2 specific risk classes require CECP
CECP not required in case any of 3 circumstances fulfilled
Expert panels
Decide
Step 1: Screening experts decide based on three criteria
Step 2: Thematic panel develops opinion
Opinion for Notified Body
21 days
60 days
CECP

## Slide 6: Notified Body Decision: which dossiers need to undergo the CECP ?
Risk class requirements
Class III implantable devices
Class IIb devices for administering and/or removing medicinal products (‘ARMP’ devices)
Circumstances under which dossiers are exempt from CECP:
Exemption 1: Renewal of a certificate issued under the Medical Devices Regulation - MDR
Exemption 2: Modified devices, subject to an assessment of adverse changes of the risk-benefit determination
Exemption 3: Common Specifications for device type exist and specific requirements are fulfilled
Exemptions 1 and 3 will, in the beginning, not be relevant: MDR becomes applicable only in May 2021 and currently there are no Common Specifications.
Exemption 2 will not be relevant initially for modifications of MDR products. However also products marketed under the Directives (‘legacy products’) may fall under exemption 2.

## Slide 7: Legacy devices and exemption 2 regarding CECP
Exemption 2: Modified devices, subject to an assessment of adverse changes of the risk-benefit determination
‘Legacy devices’ (already marketed under the Directives) fall under this exemption if the modification is limited to changes needed in order to comply with MDR requirements, for instance new requirements concerning instructions for use (IFU) documentation.
If the modification goes beyond above, legacy devices require a CECP – independent of considerations of benefit risk ratio
‘Legacy device’
Benefit-risk ratio adversely affected?
Modification limited to changes to comply with MDR ?
CECP not required
YES = exemption 2 applicable
NO
YES
NO
CECP required
Modifications go beyond those required for mere compliance with new MDR requirements

## Slide 8: Purpose of the CECP
Medical Device Regulation 2017/745 Recital 56:
“The consultation of expert panels in relation to the clinical evaluation should lead to a harmonised evaluation of high-risk medical devices by sharing expertise on clinical aspects and developing common specifications on categories of devices that have undergone that consultation process.”

## Slide 9: Scope of Expert Panel review
Out of scope:
Not a review of manufacturers’ information, although such information is taken into account when developing the opinion on the CEAR of the NB
Not a comprehensive review of conformity assessment of the device
Not a judgement of whether a certificate of conformity should be issued
Not a judgement of the overall safety / performance of the device
Instead: A review focused on the clinical evaluation assessment report and specific aspects thereof
In scope:
Advice to Notified Bodies: collegial review of Clinical Evaluation Assessment Report (CEAR)

## Slide 10: What do experts review?
Clinical Evaluation Plan
Post-market clinical follow up plan and report (where available)
Clinical Evaluation Report
Manufacturer: clinical evaluation and relevant documents
Notified Body: clinical evaluation assessment
Clinical Evaluation Assessment Report (CEAR)
Expert panels: Opinion on Notified Body’s CEAR on the basis of manufacturer’s documents
Opinion on CEAR in particular concerning
Benefit-risk determination
Consistency of clinical evidence with
  intended purpose and
  post-market clinical follow-up plan and
  Post-market clinical follow-up report (where available)

## Slide 11: The three decision criteria
Novelty of the device or of the related clinical procedure involved, and the possible major clinical or health impact thereof
Significantly adverse changes in the benefit-risk profile due to scientifically valid health concerns in respect to components, source material or impact in case of failure of the device.
Significantly increased rate of serious incidents
Actual device under assessment
From pre-market space (clinical investigation/evaluation)
Information from specific category or group of devices.
Typically from post-market space
Criterion 1 will be most relevant in practice
Criteria 2 and 3 are typically based on systematic evaluations of post-market surveillance information. Experts do not have such information, in particular what concerns serious incidents.
If relevant information is available to the Secretariat (e.g. from MDCG group on post-market surveillance), the Secretariat will forward it to panels
Thus, most dossier evaluations will be based on the use of criterion 1: novelty AND major clinical/health impact.

## Slide 12
1
2
3
4
5
Background and introduction
Criterion 3: significantly increased incidents
Criterion 1: novelty and resulting impacts
Criterion 2: scientifically valid health concerns
Overall assessment and available tools

## Slide 13: Why is novelty so important in the context of this consultation?
Review is intended to improve clinical evaluation and thus enhance ultimately patient safety in regard to high-risk devices that remain in the body for prolonged periods or which deliver/remove medicinal products
Typically novelty means lack of experience regarding safety and performance or of specific features of the device or related clinical procedure
For new products there is insufficient experience regarding future real-world safety and performance
In case of modifications of previous device variants, consider relevant real-world experiences from post-market space, e.g. in post-market clinical follow-up report

## Slide 14: Assessment of novelty and resulting clinical/health impact
Assess whether there is novelty of the device or the related clinical procedure
If there is novelty, reflect whether it may have any possible major clinical and/or health impacts
Determine degree of novelty and severity of clinical and/or health impact
Criterion 1
Yes
Yes
Determine degree of novelty & severity of clinical/health impact
Clinical/health impact of novelty?
Novelty?

## Slide 15: Novelty and resulting clinical/health impact
Criterion 1 contains a logical conjunction (‘AND’…’THEREOF’): Novelty alone or impact alone are insufficient to prompt a scientific opinion.
Health impact
Novelty
Clinical impact
Venn diagram
Opinion possibly required
‘The kidney of decision’

## Slide 16: What are possible ‘dimensions’ of novelty?
Clinical or surgical procedure related to novelty
Device-related
novelty
  Mode of use or treatment option
  Device-patient interface
  Interaction and control
  Deployment methods
Medical purpose
Design
Mechanism of action
Materials
Site of application
Components
Manufacturing process

## Slide 17: Clinical impact and Health impact
Benefits, hazards/ risks at individual level, typically based on clinical evaluation (pre-market space)
Net potential benefits & hazards/ risks on population level which may occur during post-market use (real-world conditions)
Clinical impact
Health impact
Emphasis on negative effects (hazards/ risks)
Consider also possible effects
Consider the severity of effects

## Slide 18: Assessment Matrix (optional tool)
Severity of possible impact
Novelty
dimensions

## Slide 19: Uncertainties and decision making
Novel devices will have no real world data but mainly data from optimised conditions
That means that ‘possible impacts’ may need to be estimated by experts based on to their own clinical experience. This involves uncertainties.
These need to be considered when making a decision
Uncertainties can also tip the balance towards a positive decision: for instance a high level of uncertainty may justify the need of an opinion, even if the impact severity does not appear major based on the data at hand
?!

## Slide 20: Decision Matrix (optional tool)
Need for an opinion may be justified due to a high level of uncertainty

## Slide 21
1
2
3
4
5
Background and introduction
Criterion 3: significantly increased incidents
Criterion 1: novelty and resulting impacts
Criterion 2: scientifically valid health concerns
Overall assessment and available tools

## Slide 22: Criterion 2: Scientifically valid health concerns
Scientifically valid health concerns in regard to components, source materials or health impacts in case of device failure resulting in a significant adverse change in the benefit-risk profile of a specific category / group of devices
Criterion 2
Information available?
Information reliable?
Relevant for device under assessment?

## Slide 23: Information on health concerns
If the Secretariat is in a position to forward information on health concerns (e.g. based on analyses of the MDCG group on post-market surveillance), this information may be based on systematic and careful analyses of the following data sources:
Post-market surveillance and vigilance activities of manufacturers:
  Reporting of serious incidents individually
  Periodic summary reports
  Field safety corrective actions
Periodic summary update reports by manufacturers, depending on class
Trend reports by manufacturers
Summary of safety and clinical performance by manufacturers

## Slide 24: Health concerns based on knowledge and expertise of panel
In exceptional cases experts may have information on health concerns in relation to a specific group or category of devices (e.g. relevant scientific and clinical literature)
Both reliability and relevance of such information for the device under assessment should be taken into account:
  Reliability: for example data quality, case numbers, robustness of evidence * Anecdotal evidence is not acceptable * Information relating to a product of only one manufacturer would typically not be considered sufficient to fulfil the requirement of ”specific group or category of devices” * Evidence should be traceable and documented in high-quality peer reviewed publications (e.g. conference abstracts are insufficient)
  Relevance: for example if the device under assessment is using the same or similar source materials and/or components as those used for the group or category of devices and where their use has led to health concerns
Need for a justification on reliability and relevance

## Slide 25
1
2
3
4
5
Background and introduction
Criterion 3: significantly increased incidents
Criterion 1: novelty and resulting impacts
Criterion 2: scientifically valid health concerns
Overall assessment and available tools

## Slide 26: Criterion 3: significantly increased rate of serious incidents
Concerns exclusively information provided from the Secretariat: where available, the Secretariat will provide information on significant increases of serious incidents of a specific category / group of devices
Such information does not necessarily trigger an opinion
The panel should carefully consider on a case-by-case basis if the serious incident related information is relevant for the device under assessment
Criterion 3
Information available?
Relevant for device under assessment?

## Slide 27
1
2
3
4
5
Background and introduction
Criterion 3: significantly increased incidents
Criterion 1: novelty and resulting impacts
Criterion 2: scientifically valid health concerns
Overall assessment and available tools

## Slide 28: Overall Assessment process
Yes
Yes
Determine degree of novelty & severity of clinical/health impact
Clinical/health impact of novelty?
Novelty?
Opinion required?
Device under assessment
Information from specific group/category of devices
Criterion 1
Simplified flowchart: ‘No’ decision exits not shown. A ‘no’ exit at any point closes the respective criterion.
Expert judgement based on clinical experience and taking uncertainties into account
Information available?
Information reliable?
Information relevant for device under assessment?
Criterion 2
Yes
Yes
Expert judgement
Information available?
Information relevant for device under assessment?
Criterion 3
Yes
Expert judgement

## Slide 29: Mandatory documentation of screening decision
Expert panel template for CECP
Part 1 Decision whether opinion required Summary table to be completed by screening experts
MANDATORY REQUIREMENT
Part 2: Scientific Opinion
Template to be completed by thematic experts

## Slide 30: Navigation guide for the screening step
Screening experts will receive an email with the following:
Navigation guide to the screening decision:
  Outlining a step-by-step approach regarding practical aspects and key elements
  Assessment Matrix: novelty versus severity of impact
  Decision Matrix: likelihood of opinion requirement
Commission Guidance Document
Template for expert decision and opinion, with administrative details and two parts to be completed by experts:
  Screening experts: summary table for documenting the assessment of the three criteria
  In case an opinion is required, the template will be passed on to the thematic panels, informing about the reasons as to why an opinion is required.
  Experts in the thematic panel will then complete the second part (scientific opinion)

## Slide 31: Thank you
© European Union 2020
Unless otherwise noted the reuse of this presentation is authorised under the CC BY 4.0 license. For any use or reproduction of elements that are not owned by the EU, permission may need to be sought directly from the respective right holders.
Icons from Freepik.com